# Manual of Petroleum Measurement Standards

# Chapter 7—Temperature Determination

Section 4—Dynamic Temperature Measurement

## SECOND EDITION, JANUARY 2018

### Special Notes

Neither API nor any of API’s employees, subcontractors, consultants, or other assigns make any warranty or representation, either express or implied, with respect to the accuracy, completeness, or utility of the information contained herein, or assume any liability or responsibility for any use, or the results of such use, of any information or process disclosed in this publication, or represent that its use would not infringe upon privately owned rights.

Classified areas may vary depending on the location, conditions, equipment, and substances involved in any given situation. Users of this Standard should consult with the appropriate authorities having jurisdiction.

Users of this Standard should not rely exclusively on the information contained in this document. Sound business, scientific, engineering, and safety judgment should be used in employing the information contained herein.

Users of Instructions should not rely exclusively on the information contained in this document. Sound business, scientific, engineering, and safety judgment should be used in employing the information contained herein.

Where applicable, authorities having jurisdiction should be consulted.

API is not undertaking to meet the duties of employers, manufacturers, or suppliers to warn and properly train and equip their employees, and others exposed, concerning health and safety risks and precautions, nor undertaking their obligations to comply with authorities having jurisdiction.

Information concerning safety and health risks and proper precautions with respect to particular materials and conditions should be obtained from the employer, the manufacturer or supplier of that material, or the material safety data sheet.

Where applicable, authorities having jurisdiction should be consulted.

Work sites and equipment operations may differ. Users are solely responsible for assessing their specific equipment and premises in determining the appropriateness of applying the Instructions. At all times users should employ sound business, scientific, engineering, and judgment safety when using this Standard.

All rights reserved. No part of this work may be reproduced, translated, stored in a retrieval system, or transmitted by any means, electronic, mechanical, photocopying, recording, or otherwise, without prior written permission from the publisher. Contact the Publisher, API Publishing Services, 1220 L Street, NW, Washington, DC 20005.

Copyright © 2018 American Petroleum Institute

### Foreword

Nothing contained in any API publication is to be construed as granting any right, by implication or otherwise, for the manufacture, sale, or use of any method, apparatus, or product covered by letters patent. Neither should anything contained in the publication be construed as insuring anyone against liability for infringement of letters patent.

The verbal forms used to express the provisions in this document are as follows.

Shall: As used in a standard, “shall” denotes a minimum requirement in order to conform to the standard.

Should: As used in a standard, “should” denotes a recommendation or that which is advised but not required in order to conform to the standard.

May: As used in a standard, “may” denotes a course of action permissible within the limits of a standard. Can: As used in a standard, “can” denotes a statement of possibility or capability.

This document was produced under API standardization procedures that ensure appropriate notification and participation in the developmental process and is designated as an API standard. Questions concerning the interpretation of the content of this publication or comments and questions concerning the procedures under which this publication was developed should be directed in writing to the Director of Standards, American Petroleum Institute, 1220 L Street, NW, Washington, DC 20005. Requests for permission to reproduce or translate all or any part of the material published herein should also be addressed to the director.

Generally, API standards are reviewed and revised, reaffirmed, or withdrawn at least every five years. A one-time extension of up to two years may be added to this review cycle. Status of the publication can be ascertained from the API Standards Department, telephone (202) 682-8000. A catalog of API publications and materials is published annually by API, 1220 L Street, NW, Washington, DC 20005.

Suggested revisions are invited and should be submitted to the Standards Department, API, 1220 L Street, NW, Washington, DC 20005, standards@api.org.

### Contents

Table of contents to be populated during the editorial review prior to publication.

Page

v

### Introduction

The purpose of this standard is to describe the measurement practice to accurately determine the temperature of liquid hydrocarbons under dynamic (i.e. flowing) conditions. The accurate measurement of temperature is essential to accurately determine the average temperature of the hydrocarbons liquid transferred, for the purpose of correcting volumes to standard conditions, computing density from equations of state, converting measured density to alternate conditions, and compensating fluid and equipment for temperature effects.

vi

## Chapter 7—Temperature Determination Section 4—Dynamic Temperature Measurement

### Scope

This standard describes methods, equipment, installation, and operating procedures for the proper determination of the temperature of liquids and gases under dynamic conditions in custody transfer, fiscal, allocation, and inventory applications. This standard describes the accuracy requirement and the calibration of the temperature measurement equipment.

This standard does not cover dynamic temperature measurements of refrigerated or cryogenic liquids.

### Normative References

The following standards and other industry publications contain provisions that, through reference in this text, constitute provisions of this standard. All standards are subject to revision, and parties to agreements based on this standard are encouraged to use the most recent editions of the standards indicated below:

API Manual of Petroleum Measurement Standards (MPMS)

- Chapter 1, Terms and Definitions, http://chapter1.api.org

- Chapter 4 (All Relevant Sections), Proving Systems

- Chapter 5 (All Relevant Sections), Metering Systems

- Chapter 6 (All Relevant Sections), Metering Assemblies

- Chapter 7.1, Liquid in Glass Thermometers

- Chapter 12.2, Calculation of Liquid Petroleum Quantities Using Dynamic Measurement Methods

- Chapter 21.2, Electronic Liquid Volume Measurement Using Positive Displacement and Turbine Meters

API Recommended Practice 500, Classification of Locations for Electrical Installations at Petroleum Facilities Classified as Class I Division 1 and Division 2

API Recommended Practice 2003, Protection Against Ignitions Arising Out of Static, Lighting, and Stray Currents

##### ASME

ASME PTC 19.3TW_2016, American National Standard for Thermowells

##### ASTM

- ASTM E1, Standard Specification for ASTM Thermometers

- ASTM E77, Standard Test Method for Inspection and Verification of Liquid-in-glass Thermometers

- ASTM E344, Terminology Relating to Thermometry and Hydrometry

- ASTM E644, Test Methods for Testing Industrial Resistance Thermometers

- ASTM E1137, Standard Specification for Industrial Platinum Resistance Thermometers

- ASTM 2251, Standard Specification for Liquid-in-Glass ASTM Thermometers with Low-Hazard Precision Liquids

- ASTM E2593, Standard Guide for Accuracy Verification of Industrial Platinum Resistance Thermometers

- ASTM E2877, Standard Guide for Digital Thermometers

- ASTM E2488 Guide for the Preparation and Evaluation of Liquid Baths Used for Temperature Calibration by Comparison

- ASTM E3186 Guide for Use and Testing of Dry-Block Temperature Calibrators

##### IACS

E10, Test Specification for Type Approval

##### NFPA

National Electrical Code

##### NIST

NIST Special Publication 250-22, Platinum Resistance Thermometer Calibrations

### Terms and Definitions

##### 3.1

Ambient conditions

The conditions (pressure, temperature, humidity, etc.) of the medium surrounding an object such as the case of a meter, instrument, transducer, etc.

##### 3.2

##### digital thermometer (DCT) (referred to as Digital Thermometers)

An electronic device (per ASTM E2877) consisting of a digital display and associated temperature sensing probe.

##### 3.3

Partial Immersion Thermometer

A thermometer designed to indicate temperatures correctly when the bulb and a specified part of the stem are exposed to the temperature being measured.

##### 3.4

Resistance Temperature Element

An electrical temperature-sensing element in common use to measure the temperature of the contents of a storage tank or the contents of a pipeline.

##### 3.5

resistance temperature detector (RTD)

A temperature measuring device that operates on the principle of a change in electrical resistance in wire as a function of temperature.

##### 3.6

Temperature Element

One or more sensors housed in a casing such as a metal probe or protecting tube that makes up a single temperature measuring unit for a thermometer.

##### 3.7

Temperature Measurement System

Consists of a sensor, transmission medium, and readout equipment and is the operating configuration used to determine the temperature of a liquid for measurement purposes.

##### 3.8

temperature sensor

A temperature sensor is one part of a thermometer in which some physical change occurs with temperature and converts this change into a value on a scale, (e.g. the scale on a liquid-in-glass thermometer, a digital device that displays a unit of measure).

##### 3.9

temperature transmitter

A device that converts the temperature measured by the sensor(s) to electrical or electronic signal and transmits the signal to a remote or local display.

##### 3.10

Thermistor

A temperature-sensitive element consisting of a semiconductor such as germanium oxide, whose conductance changes with its temperature. Thermistors are used to make temperature measurements electrically.

##### 3.11

Thermocouple

A thermocouple is a junction between two different metals that produces a voltage related to temperature difference.

##### 3.12

Thermometer

A device that measures temperature using any of a variety of different principles. A thermometer has two important components: a temperature sensor in which some physical change occurs that is dependent on temperature, and some means of indicating and/or transmitting this physical change as a value

##### 3.13

Thermometer Response time

The time required for a thermometer to indicate 63.2 percent of the magnitude of a change in the measured temperature.

##### 3.14

Thermowell

Pressure and liquid tight receptacle adapted to receive a temperature sensing element and provided with external threads, flanges or other means for pressure tight attachment to a vessel. A thermowell allows the temperature sensor to be removed and replaced without compromising the process.

##### 3.15

Total Immersion Thermometer

A liquid-in-glass thermometer designed to indicate temperatures correctly when just that portion of the thermometer containing the liquid is exposed in the temperature being measured.

Thermal radiation

Heat conduction

### Acronyms

API	American Petroleum Institute

ASTM	American Society for Testing and Materials

FCC	 Federal Communications Commission FM	Frequency Modulation

IACS	International Association of Classification Societies IS	Intrinsically Safe

IEC	International Electrotechnical Commission  Commission IMO	 International Maritime Organization

MPMS	Manual of Petroleum Measurement Standards

NEC	National Electric Code

NIST	National Institute of Standards and Technology NFPA	National Fire Protection Association

NRTL	Nationally recognized testing laboratory

DCT	 Digital Contact Thermometer DT	Digital Thermometer

RF	Radio Frequency

RFI	Radio Frequency Interference RP	 Recommended Practice

RTD	Resistance Temperature Device

SCADA	Supervisory Control and Data Acquisition UL	 Underwriters Laboratories

### General Precautions

### Introduction

Temperature can have a significant effect on the accurate determination of fluid quantities and when correcting to standard conditions. Temperature determination shall be consistent with custody transfer practices and meet the desired measurement uncertainty.

### Safety

This standard may involve hazardous materials, operations, and equipment. It is the responsibility of the user of this standard to establish appropriate safety and health practices and determine the applicability of regulatory limitations prior to use.

### Equipment Precautions

### Equipment and Design Requirements

### General

The following parameters shall be addressed when selecting a temperature device to be used for temperature determination under dynamic conditions:

- Accuracy requirements

- Response time

- Temperature measurement range of operation

- Ambient conditions

- Resolution of the temperature device to meet the discrimination for the intended application

- Safety of operation

- Compatibility with substances being measured

The selection of dynamic temperature measurement equipment and its design requirements depend on the application. Example applications include its use in prover calibration, meter proving, or calculation of the gross standard volume.

All dynamic temperature devices shall be designed and operated in compliance with safety regulations, API and relevant national measurement standards and classification rules including region specific electrical codes and regionally recognized testing laboratory certification.

All temperature equipment shall be suited for the electrical classification for its location of use. They shall conform to applicable sections of the National (UL, FM, FCC, NEC, etc.) electrical safety standards or International (IMO, IEC, IACS, CENELEC, ISO, etc.) electrical safety standards. All equipment shall be maintained in a safe operating condition and manufacturers’ maintenance instructions should be complied with.

All of the temperature equipment shall be capable of withstanding the design conditions for pressure, temperature, operation, and the environment.

All equipment shall be designed for electromagnetic compatibility (EMC), complying with user requirements and other proper standards, which means that the equipment shall neither interfere with nor be affected by interference from other equipment. Refer to IACS Unified Requirements E10—Test Specification for Type Approval for specific tests and requirements.

The equipment shall be designed with the capability to operate under variation in power source voltage according to class environmental requirements, e.g. IACS Unified Requirements E10—Test Specification for Type Approval.

### Minimum Acceptable Resolution

Table 1 provides a guide to the minimum acceptable resolution of the display or readout for dynamic temperature determination in custody transfer measurement of petroleum and petroleum products. For digital displays, the resolution is the last digit displayed. For analog displays, the resolution is the separation between two consecutive division marks on the scale without discriminating a measurement in between scale division marks. If the glass thermometer temperature is shown between two divisions, it shall be read and reported to the nearest half division.

The minimum resolution of the display may not be the same as the discrimination level. The required discrimination level, which is dependent on application, is defined in API MPMS Chapters 4 and 12. A temperature device with a higher resolution is not necessarily an indication of a higher accuracy. The use of a temperature device that can provide a higher resolution than that outlined in Table 1 is acceptable and the use of temperatures obtained may be used.

Table 1—Minimum Resolution of Reading

|  | Temperature Device Type |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Application | Liquid-in-Glass Thermometers (Graduation)b |  | Digital Thermometer |  | Temperature Transmitter (with temperature sensor) |  | RTD (fixed, with a local display only) |  |
|  | °C | °F | °C | °F | °C | °F | °C | °F |
| Meter Ticket | 0.2 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 |
| Meter Proving | 0.2 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 |
| Prover Calibration | 0.2 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 |
| Density Meter Proving | a | a | a | a | a | a | a | a |
| Pynometer Calibration | a | a | a | a | a | a | a | a |
| Density Meter Calibration | a | a | a | a | a | a | a | a |
| a Refer to API MPMS Chapter 9.4 for reference temperature values. b Non-ASTM liquid-in-glass thermometers with higher resolutions are also available. |  |  |  |  |  |  |  |  |

### Thermowell

Thermowells are classified based on their use as test wells and sensor thermowells as explained in 6.3.1 and 6.3.2.

Parameters that affect heat transfer between the fluid and thermowell include;

Fluid velocity

  Higher velocity increases convective heat transfer between the fluid and the thermowell.

  This improves the thermometer’s response time and accuracy.

  Low velocity may lead to poor thermal contact and slower response.

Fluid temperature

  The temperature gradient between the fluid and the thermowell drives heat transfer.

  A larger gradient results in faster heat exchange but may also cause thermal stresses in the thermowell material.

Flow Regime

  Turbulent flow enhances mixing and increases heat transfer coefficients.

  Laminar flow has lower heat transfer efficiency due to limited mixing and lower velocity near the thermowell surface.

Thermowell Design and Material

  Material conductivity affects how quickly heat is transferred from the fluid to the thermometer.

  Wall thickness and insertion length influence thermal lag and mechanical strength.

  A longer insertion into the flow improves thermal contact but may increase vibration risk.

Sensor contact with thermowell bottom / heat transfer fluid. - Add

Pipe pressure

  Pressure affects fluid density and can influence the heat transfer coefficient.

  In gases, higher pressure increases density, improving heat transfer.

  In liquids, the effect is less pronounced but still relevant in compressible flow scenarios.

Ambient Conditions

  External pipe temperature or insulation can affect the heat loss or gain from the system, impacting thermometer readings.

Thermometer type

  RTDs, thermocouples, and thermistors have different response times and sensitivities.

  Their placement inside the thermowell and thermal contact quality matter.

Thermowel l response time estimation

Estimating the response time of a thermometer in a thermowell (often called "thermal response time" or "time constant") involves understanding how quickly the thermometer reacts to changes in fluid temperature. The response time depends on:

- Thermowell material and thickness

- Thermometer type

- Fluid velocity and properties

- Heat transfer coefficient

ANNEX D provides a description of the lumped capacitance model that can be used to estimate the response time of a thermometer in a thermowell.

ANNEE E provides an example calculation.

### Thermowell Type

      - Test Thermowell

Test wells are thermowells installed for occasional use, (temperature device verification, calibration and periodic fluid temperature measurement) and capped when not in use. Test wells shall be installed as close as practical to the sensor wells. A suggested maximum distance between the test well and sensor thermowell is three pipe diameters or 457 mm (18 in.) whichever is less. The test well should ideally match the sensor well thickness and construction materials and shall have the same immersion depth and a bore designed to accommodate the insertion of the verification temperature device. Capping prevents foreign material from accumulating in the wellbore. A clogged thermowell may cause measurement errors and may damage thermometers. It is recommended that thermowells be checked periodically for an accumulation of foreign material and cleaned if necessary.

Operating conditions influence how quickly that a the thermowell will represent the temperature of the fluid.  Consult with the Thermowell manufacturer on design of a thermowell to ensure heat conducting properties of the thermowell are acceptable.

##### Sensor Thermowell

Sensor wells are thermowells installed for use with a permanently installed temperature device, and shall be matched to the temperature device.

##### Selection of Thermowell

Thermowells shall be selected based on the application criteria described 6.3.3, 6.3.4, and 6.3.5. Calculations should be performed in accordance with ASME PTC 19.3 TW to confirm the appropriateness of each selected thermowell to resist flow-induced vibration.

mechanical strength requirements take precedence over measurement accuracy.

Excerpt from PRCI Report  CFD Analysis of the Heat Transfer Characteristics and the Effect of

Thermowells Used in Natural Gas Measurement Facilities

Recommendations

(1).The stepped thermowell is recommended among the three types of thermowells as it showed the smallest error for both small and large pipe installations.

(2).For the same type of thermowell, a larger pipe installation can result in smaller temperature measurement error.

(3).The tapered thermowell is not recommended for the small pipe installation as it had the largest error among the three thermowells. In the large pipe installation, the volume of the tapered and the straight thermowell is important. It is a good practice to measure and calculate the volume of the thermowells in order to select the proper thermowell. The tapered thermowell is preferred if its volume is smaller than the straight thermowell.

(4).Further research on the following topics can be valuable to the members: (i) how do the finned and the taper/step thermowells work and (ii) how do the glycol, thread sealant and Teflon tape affect the heat transfer and the accuracy of temperature measurement.

(5).Although the flange thermowell installation is not as common as the threadolet installation in the industry, it showed a smaller temperature measurement error than the threadolet installation based on the preliminary CFD simulations. It is worthy of further study.

(6).More data points will be needed to obtain correction factors between the temperature error and the flow rate because the relationship between the effect on the temperature measurement and the combination of the pipe diameters, types of thermowells, thermal conditions at the pipe wall surface, and various gas velocities is complex.

Review API RP 551

##### Pressure Rating

The thermowell selected shall comply with design codes for the operating pressures and temperatures of the system.

##### Installation

The thermowell should be selected to conform to code and user installation practices. Thermowells are typically threaded, welded, or flange mounted. The immersion length of the thermowell should be sufficient to put the sensing portion of the device within the center one-third of the pipe diameter or provide immersion of 0.3 m (1 ft.) unless limited by fluid velocity considerations. The thermowell should be installed in as near to a vertical position as practical to allow it to be filled with an appropriate thermal-conducting material to cover the sensor. . Thermowells should be installed in the same pipe size as the meter piping except for small diameter piping where the Thermowell can block fluid flow. In this case, the thermowells should be installed in piping at least one nominal pipe size larger than the meter. Test thermowells should have an internal bore sized from 0.375” to 0.500” to allow for insertion of a reference thermometer. Test thermowells should have a minimum internal depth of 3” to allow for full immersion of the reference thermometer. Test thermowells should have a minimum thermometer insertion clearance that allows the unobstructed insertion of a reference thermometer.

Some flow measurement devices (ultrasonic, orifice, and turbine meters) require that the thermowells be installed outside the flow meter downstream conditioning area, see chapter 5 and chapter 14 (orifice) for specific requirements. For turbine meters, the temperature device shall be installed downstream of the meter at least 3D downstream of the meter and outside the flow conditioning area.

For location of temperature devices and thermowells on provers, see chapter 4.

| Meter Type | Position | Minimum D | Maximum D |  |
| --- | --- | --- | --- | --- |
| Orifice | Downstream | 5 | 15 |  |

##### Material

The thermowell should be constructed of a material that is compatible with the liquid material that the thermowell is exposed to and should provide a degree of corrosion resistance for all surfaces. Usually Type 304 or 316 stainless steel is specified.

##### Thermowell Conductive Material

In the case where the temperature sensor does not come into contact with the thermowell wall, the space between the sensor and the thermowell wall shall be filled with an appropriate amount of non-freezing thermal-conducting material to improve heat conduction between the wall of the thermowell and the sensor, and to improve the temperature sensor response time.

##### Electrical Temperature Sensor and Transmitter

      - Temperature Sensors

Electronic temperature devices for dynamic measurement are usually housed in metal probes that mount into thermowells. The probes are generally tip sensitive. Therefore, the probes must be securely seated in the bottom of the thermowell for optimum temperature transfer. Spring loaded or adjustable length probes are recommended. If the probe cannot be seated in the bottom of the thermowell, an appropriate thermal-conducting material shall be used between the temperature sensor and the thermowell wall.

The wiring to the probe is critical because of the low signal levels of the devices. These devices should be installed as recommended by the manufacturer for best accuracy. The equipment and transducers should be installed in accordance with all API and NFPA National Electrical Code (NEC) hazardous area specifications.

##### Resistance Temperature Detector

A resistance temperature detector (RTD) is a sensor with an electrical resistance that is a function of temperature. The RTD is usually a small coil of platinum (or copper wire) and when used with appropriate circuits will provide temperature signals to readouts and other equipment.

The RTD may be a resistance wire wound on a supporting nonconductive core, a thin film type, or other type. The element should be properly encased in a protective enclosure (probe), typically constructed of stainless steel. The electronic circuits should be intrinsically safe as required. The length of the temperature sensitive portion of a single- point (spot) element should not exceed 100 mm (4 in.). Three-wire or four-wire RTDs are recommended to compensate for lead length resistance.

Refer to ASTM E1137 for material and fabrication specifications and ASTM E644 for manufacturer calibration and testing of RTDs.

Class A RTDs are recommended for accurate temperature measurement such as custody transfer. For such applications, accuracy may be improved by reducing the span of the transmitting device to under 40 °C (100 °F) from the manufacturer wide operating range where appropriate.

##### Thermistor

Thermistors may be used as an alternative to the resistance temperature device for digital thermometers, but are not used as commonly as an in-line temperature sensor. Given the small size of the semiconductor used in the sensing portion of a thermistor, the temperature sensor will usually have a very quick response. With a suitable choice of the semiconductor, linearity of response can be obtained over a wide range of temperatures. Some drift may be experienced with thermistors, and errors can arise from auto-heating of the semiconductor. Adjustment facilities are therefore usually incorporated in thermistor-based digital thermometers. Thermistors are often not stable and accurate enough for custody transfer purposes.

##### Other Types of Electrical Temperature Sensors

Other types of temperature sensors (e.g. thermocouples, semiconductors, etc.) are also available and may be suitable for custody transfer purposes, provided that they meet the performance requirements in this standard.

##### Temperature Transmitter With and Without Temperature Sensor

      - General

A temperature transmitter is a device that converts a signal from a temperature sensor into a form suitable for propagating the temperature data by the sensor from the site of measurement to the location where the data will be displayed and used. Transmitters may be located remotely from the sensor or directly attached as a single assembly.

##### Electrical Design Safety

Electronic transmitters shall meet the requirements for the National Electrical Code (NEC) hazardous area where they are installed. The minimum shall be explosion proof, Class I, Division 2, Group D.

When used in Intrinsically Safe (IS) systems, transmitters shall be specified as IS and shall have an IS certification label. When an IS instrument is utilized, special wiring practices are required.

##### Location

Mounting the temperature transmitter as close as practical to the sensor, including using transmitter and sensor single assemblies, eliminates the need for special wire, thermocouple signal wire segregation and specialized low-voltage inputs on devices receiving the temperature signal. Utilize an acceptable signal format or communication protocol, for example, the 4 to 20 mA output to the remote readout (e.g. flow computer, temperature averager, etc.).

##### Electronic Temperature Transmitter Design and Installation Requirements

Recommended design and installation requirements for temperature transmitters are described below:

- Intelligent transmitters use digital data processing and communication methods for performing its functions and for safeguarding and communicating data and information on its operation. It may be equipped with additional sensors and functionality which support the main function of the intelligent transmitter. The variety of added functionality can for instance enhance accuracy and rangeability, self-test capabilities, and alarm and condition monitoring.

- RTD transmitters shall use resistance bridge-type temperature measuring circuits. The housing of the temperature transmitters shall have dual compartments (or equivalent) to separate the electronics from the wiring connections, which shall be in an integral or attached junction box. RTD transmitters shall use reference-based measuring circuits. RTD transmitters shall compartmentalize the wiring connections independently from the electronics. The compartments will permit access to cable connections without exposing the electronics. The compartments can join with each other through a cable gland, a module connector, or equivalent.

- Temperature transmitters shall be protected from radio frequency interference (RFI) and electromagnetic and electrostatic interference. They shall include adjustable damping to smooth out transient signal noise to the extent that it will not compromise accuracy requirements of teh device

- Range and span shall be adjustable without changing electronic components. Inputs and outputs should be isolated. Failure of the sensor shall result in an alarm or predicable fail-safe action (commonly referred to as burnout protection). Commonly, loss of input results in the transmitter output going to the maximum upscale value of its range.

### Digital Thermometer

Digital Thermometers designed for dynamic measurement shall meet ASTM E2877 and the accuracy requirements specified in Section 9. The accuracy required depends on its application (i.e. whether it is used as a calibration reference or as a working device).

### Liquid-in-Glass Thermometer

      - General

Liquid-in-glass thermometers designed for dynamic measurement shall meet the accuracy requirements specified in ASTM E1 or ASTM E2251. Their use and their accuracy required depends on its application (i.e. whether it is used as a calibration reference or as a working device).

##### Partial-immersion and Total-immersion Liquid-in-Glass Thermometers

Usage at different immersion depths or at significantly different ambient temperatures from which a thermometer was certified may require that stem corrections be made. Total-immersion thermometers are typically used partially immersed into a thermowell. The portion of the stem that is outside the thermowell will respond to the ambient temperature, and may cause a significant error in the reading, depending on the difference between the temperature being measured and the ambient temperature and the amount of active stem that is exposed to the ambient condition.

Corrections should be made where necessary, in accordance with the procedure in API MPMS Chapter 7.1, Liquid-in- Glass Thermometers, or ASTM E77, Standard Test Method for Inspection and Verification of Thermometers.

##### Permanently-installed Glass Thermometer

Permanently-installed, liquid-in-glass thermometers used for custody transfer shall meet the accuracy requirements specified in Section 9. The accuracy required depends on its application (i.e. whether it is used as a calibration reference or as a working device).

### Dial Thermometer

Dial thermometers shall not be used for temperature determination in fiscal and custody transfer applications.

### Local and Remote Readout

Depending on the nature of the application the temperature reading may be displayed locally using a weather proof display or at remote site. In some applications both are required. The reading from both the remote and the local display should be identical to the required discrimination level. Remote transmission of signals over a long distance may degrade the reliability and accuracy.

### Installation Requirements

### General

The purpose of dynamic temperature determination is to determine the temperature of the fluid and by inference, the temperature of the metering device, meter or prover, in order to calculate fluid and equipment temperature correction factors. Proper installation of temperature determination equipment is necessary for accurate temperature determination without affecting meter performance.

### Meters

##### General

Barring guidance to the contrary in API MPMS Chapters 5 and 6, the temperature device should be located as close as possible to the meter, downstream and preferably no further than five pipe diameters from the meter run. Bidirectional meters may require some compromise, with the temperature measurement device or devices located outside the bounds of flow-conditioning elements.

Where several meters are manifolded in parallel, one temperature device located in the total liquid stream is acceptable, provided the temperature at each meter and at the temperature device location agrees within the accuracy limits of 0.25 °C (0.5 °F), for volume measurement of fluids with vapor pressure below 101.325 kPa (14.696 psia). Test thermowells are provided for each meter run to verify that the individual meter-run temperatures compare with that of the total liquid stream. This verification should be done as frequently as necessary to satisfy all parties involved.

##### Displacement Meter

The temperature device for a displacement meter shall be installed in the meter body or in close proximity to the meter body, preferably no more than five pipe diameters and downstream of the meter when practical. Installation on the meter body will typically require installations that differ from the recommendations given in 6.3.1 for the test wells and sensors. Refer to API MPMS Chapter 6.

##### Turbine Meter

The temperature device should be installed downstream of the meter. To prevent flow profile distortion, sensor wells shall be located close to the meter but not to be less than five pipe diameters from the outlet flange of the meter. This would include the sensor well and test well since these devices disrupt flow patterns. In case of a bidirectional meter the temperature devices shall be mounted outside the boundaries of flow-conditioning elements or piping. In general, the downstream location is preferred since it will not affect flow-conditioning requirements.

Caution should be exercised if placing the temperature device upstream in a turbine meter run so as not to disturb the flow profile. For more information on the typical installation for turbine meter, refer to API MPMS Chapters 5.3 and 6.

##### Ultrasonic Flow Meter

Temperature sensor installed for a metering system using ultrasonic flow meter (USM) shall be installed in accordance with the requirement of API MPMS Chapters 5.8 and 6 or the manufacturer installation requirements.

The thermowell, temperature sensor and transmitter shall be installed to minimize any possible noise affecting the output signal of the USM transducer.

##### Coriolis Force Mass Flow Meter

The temperature device in a Coriolis force mass flow meter intended to measure the flow tube temperature is not designed and does not have the required accuracy to determine the fluid temperature in a custody transfer application. The mass flow meter should have accurate temperature measurement even if it is not used for volumetric measurement. The temperature measurement is needed for both the base and flowing density.

Thermowells should be installed near the flow sensor so that the measured temperature is representative of the fluid temperature in the Coriolis meter. Normal practice is to install the thermowell downstream of the meter. Installation of a test thermowell next to the sensor thermowell is recommended.

If the density is calculated based on temperature and pressure, additional pressure and temperature measurements upstream and downstream of the Coriolis meter flow sensor may be required to determine the average density in the flow sensor. Refer to API MPMS Chapters 5.6 and 6 for more information.

##### Other Flow Meters

Guidance for temperature measurement device installation for other meters is the same as that for a turbine meter unless specified otherwise by the vendor.

### Provers

##### Displacement Prover

When determining the temperature of a liquid flowing through a prover, for use in correcting for the thermal effects on the liquid and prover, the liquid temperature shall be measured as close to the prover inlet or prover outlet or both as is practical. When two measurements are taken, they shall be averaged to arrive at the prover temperature.

In many cases, the inlet and outlet liquid temperature are the same (e.g. on provers with high flow rates or on well insulated provers). Under these conditions, one temperature sensor is acceptable if temperature measurements are performed periodically on the opposite end of the prover to verify that the two temperatures agree within the tolerances given in Table 2. Refer to API MPMS Chapter 4.

##### Tank Prover

When determining the temperature of a liquid contained in a tank prover, for use in correcting for the thermal effects on the liquid and tank material, the following numbers of temperature sensors are recommended for both insulated and non-insulated tanks:

- Tanks with less than 380 L (less than 100 gal)—1 sensor.

- Tanks with 380 L to 1900 L (100 to 500 gal)—2 sensors.

- Tanks with greater than 1900 L (greater than 500 gal)—3 sensors.

If one temperature sensor is used, it should be placed in the center of the tank vertical height. If two temperature sensors are used, the readings will be taken from the middle of the upper half and the middle of the lower half of the vertical height. If three temperature sensors are used, the readings will be taken from the middle of the upper third, the middle of the middle third and the middle of the lower third of the tank vertical height. When two or more temperature sensors are used, they should be equally spaced around the tank circumference.

Where a tank prover is insulated or its cycle is rapid, two temperature sensors may be sufficient in tanks over 1900 L (500 gal) capacity. The temperature sensor shall be positioned a minimum of 0.3 m (1 ft.) from the tank shell. Refer to API MPMS Chapter 4.

##### Master Meter

The guidance for temperature measurement devices for line meters applies to the master meters of the same technology.

##### Field Standard Test Measure

No fixed thermowell is normally installed on field standard test measures. Temperature determination shall be determined, in this case, by manual temperature measurement.

### Test Facilities

A test thermowell should be provided close to temperature sensors that are permanently mounted in piping systems, such as in meter runs or meter prover systems. Refer to 6.3.1.1.

### Signal Transmission to Remote Readout

##### General

The requirements for data collection, transmission, and receiving vary with the type and the make of the temperature device and the system. For example, an electronic temperature transmitter is often part of flow metering system. The manufacturer recommendations should be followed. Additional requirements may be necessary to provide proper security and protection of the measured data. The signal cabling installation should conform to all applicable codes and regulations as referred to in Section 5, and the following sections of this document.

##### Radio Frequency Interference

Particular attention should be made to avoid radio frequency interference. Cable shielding and cable routing should be designed to minimize interference. Filtering may be required at equipment inputs.

##### Signal Wiring

Signals are typically transmitted in pairs of twisted, shielded conductors in an insulated multi-pair cable installed in conduits or buried. The line impedance should be calculated to stay within the maximum impedance specified by the equipment manufacturer. Digital signals are recommended instead of analog signals, especially if long distance transmission or high impedance is expected. Alternatively, signals may be transmitted via other media (for example, fiber optics or coaxial cables) as recommended by the equipment manufacturer.

##### Grounding

Proper grounding is important to protect equipment from damage due to transients or surges, which can result in loss of measurement data. Grounding requirements vary by type and make of the equipment. Therefore, the manufacturer recommendations should be explicitly followed. In aerial installations, the supporting messenger should be bonded to ground. In buried cable installations, only cable suitable for direct burial should be used. In conduit systems, continuity of ground through conduit joints should be ensured either by proper makeup of joints or by bonding connections around each joint.

##### Wiring Shields

The wiring shields may be copper, aluminum, or steel in accordance with the manufacturer recommendations. The overall shielding should be bonded together at all junction boxes and properly grounded on only one end either to a power line grounded neutral or to a driven ground rod.

##### Signal-to-Noise Ratio

The requirements covered in 6.3.8.4 should be followed to provide immunity to noise pickup.

##### Receiving Unit

The receiving unit may be an integral part of temperature system. The unit may be in a local configuration or in a remote configuration (tank farm control house). The remote temperature readout unit should be capable of the following:

- Display temperatures on a real-time basis.

- Perform data validity checks and alert the operator if errors are detected.

- Display alarms such as high temperature, low temperature, and so forth.

The data collection, data transmission and receiving equipment should be designed and installed (refer to API RP 500 and RP 2003) such that data transmission and receiving should:

- not compromise the accuracy of the measurement. The difference between the temperatures displayed by the remote receiving unit and displayed (or measured) by the temperature transmitter should not exceed ±0.1 °C (±0.2 °F) (with the temperature input stable during the comparison) if electronic analog transmission, and the difference should be zero if digital transmission method is used;

- not compromise the resolution of the measurement output signal;

- provide proper security and protection of the measured data to ensure its integrity;

- provide adequate speed to meet the update time required for the receiving unit to be electro-magnetically immune.

### Procedures for Temperature Determination

### General

Dynamic temperature determination may require the use of a liquid-in-glass thermometer, digital thermometer, thermistor, RTD, or temperature transmitter depending on the application. These temperature devices are acceptable provided that they meet the accuracy stated in Section 9.

### Temperature Measurement for Prover Calibration

For each prover calibration, one or more fluid temperatures shall be determined and reported on the prover calibration report. Temperature measurement during a prover calibration shall follow the procedure described in various sections of API MPMS Chapter 4.9, depending on the type of prover involved and the method used.

### Temperature Measurement for Meter Proving

For each proving run the temperature shall be determined and reported on the proving report. Once the fluid temperature in the meter run and prover has stabilized, the average temperature of each proving run shall be determined for both the prover and the meter (if applicable) independently. Where it is required to correct fluid temperature for both the meter and prover, it is not allowed to use the prover or the meter temperature for both. Refer to API MPMS Chapter 4.8.

### Temperature Measurement for Online Density Meter Proving

For each density meter proving run the temperature shall be determined and reported on the proving report. Electronic devices are preferred because of their readability. Refer to API MPMS Chapter 9.4.

To eliminate possible errors, the same device should be used at all three points (meter, density meter and proving device). The temperatures between the density meter, proving device (i.e. pycnometer), and volume meter should coincide as closely as possible, such that the following criteria are met:

- During normal operation, the density deviation between the density meter, pycnometer, and volumetric meter shall not exceed 0.05 %.

- See API MPMS Chapter 9.4, Annex A for specific fluid requirements as the acceptable error is a function of temperature, process application and fluid type.

### Temperature Measurement for Meter Ticket

Temperature used in the measurement ticket shall be the flow-weighted average temperature of the batch or parcel transferred. Where an electronic flow computer or a metering system control computer system is available, refer to API MPMS 21.2. Where there is no electronic flow computer, a manually-determined average temperature of a transfer may be used. In the event of a failure to the computing system used to calculate the measurement ticket, a manual ticket calculation shall use the temperature history as basis for the temperature reading. In custody transfer with electronic flow measurement, the average temperature of a meter ticket or bill of lading should be a flow- weighted average temperature. Refer to API MPMS Chapters 12.2 and 21.2.

### Use of Manual Measurement

A manual temperature measurement may be used if a spot temperature is required provided the accuracy and discrimination levels in Section 9 are met. The manual temperature measurement may be required under the following circumstances:

- Verifying or calibrating the working thermometer

- Calibration of flow meter by a tank prover

- Calibration of tank prover, and field standard test measure

- Inspecting or testing the functionality of a working thermometer

- Working out a manual measurement ticket in case of the failure of inline, automatic method

- Calibration of a prover by a field standard test measure

The temperature shall be allowed to stabilize before a reading is taken. The temperature shall not fluctuate by more than the range of temperature values in Table 2, for the intended application.

Before initial use, each temperature measurement device shall be compared to a reference device traceable to a national metrology institute (NMI) such as NIST.

### Accuracy Requirements

### Accuracy of Temperature Measurement Devices

##### Minimum Accuracy

Dynamic temperature measurement devices in custody transfer service shall be calibrated to demonstrate that they meet the accuracy of the intended service. Specifically they shall meet the tolerances in Table 2.

##### Table 2—Minimum Accuracy Requirements of Temperature Devices

| Commodity | Crude Oil, Refined Products, Lubricating Oil |  |
| --- | --- | --- |
| Application |  |  |
|  | °C | °F |
| Meter Ticket | ±0.25 | ±0.5 |
| Meter Proving | ±0.25 | ±0.5 |
| Prover Calibration | ±0.25 | ±0.1 |
| Density Meter Proving* | a | a |
| Pynometer Calibration* | a | a |
| Density Meter Calibration* | a | a |
| a Refer to API MPMS Chapter 9.4 for reference temperature values. |  |  |
| Commodity | Natural Gas |  |
| Application |  |  |
|  | °C | °F |
| Meter Ticket | ±0.5 | ±1.0 |
| Meter Proving | ±0.5 | ±1.0 |
| Prover Calibration | ±0.25 | ±0.1 |

Where a temperature transmitter is used, the transmitter accuracies, which include both the temperature sensor and the transmitter should be stated as plus or minus () temperature values. The stated accuracy of a transmitter can be expressed as:

- a percentage of the upper range value (URV), or

- a percentage of the calibrated span.

For example, a transmitter with a URV of 250 °F which has been calibrated over a span of 0 °F to 180 °F. Assume flowing temperature to be 100 °F. Therefore, if the transmitter accuracy stated as (shown in U.S. customary unit):

- 0.25 % of URV, then the accuracy is 0.625 °F

- 0.25 % of calibrated span, then the accuracy is 0.45 °F

It is important to understand what the manufacturer stated accuracy in the specifications mean with respect to accuracy, before choosing the equipment for your application.

The accuracy of temperature transmitters after installation and in operation can be affected by:

- Ambient temperature effects

- Vibration effects

- Power supply effects

- Mounting position effects

It is important to evaluate the effects of these conditions, and calculate their effect to determine the “as installed” accuracy of the temperature transmitter.

##### 9.1.2 Accuracy of Calibration Reference

A reference thermometer is any temperature device that is used to verify or calibrate another temperature device and shall be calibrated annually. The accuracy and calibration requirements is shown in the table below.

Table 5—Requirements for Test Equipment

| Test Equipment | Resolution | Accuracy Requirements 1 | Frequency of Calibration |
| --- | --- | --- | --- |
| Digital Thermometer | 0.1 | 2-3 times better than the device being tested | Yearly |
| Portable Electronic Thermometer (PET) | 0.1 | ±0.3 °C (±0.5 °F) See note 3 | Yearly |
| Liquid-in-glass | 0.1 | 2-3 times better than the device being tested | Yearly |
| RTD Simulator | N/A | 3 times better than the device being tested | Yearly |
| Thermal Blocks and Baths | N/A | 3 times better than the device being tested | Yearly |
| Multimeter – current measurement (4-20 mA) | 0.01 | x.xx % TBD | Yearly |

Improvements in transmitter accuracy can result in difficulty obtaining field verification/calibration equipment that is significantly more accurate than the transmitter being verified/calibrated. In some cases this may lead to a maintenance practice of using factory calibrated devices and only verifying the device to verification tolerances in the field (see 8.2).

Where possible, verification/calibration equipment should be at least twice as accurate as the device being calibrated. As a minimum standard, calibration/verification equipment shall have either an accuracy less than or equal to 0.1 % of the span/operating range of the device being calibrated or less than or equal to 0.1 % of reading. Temperature standards shall have an accuracy of less than or equal to 0.5 °F/0.3 °C. This will make the verification/calibration equipment for EGM systems reasonably available. Use of calibration equipment that is less accurate than the equipment being calibrated will add significantly to the calibration uncertainty and needs to be considered in the facility uncertainty calculation. Calibration/verification equipment with a specification that is different from the span/operating range of the device being calibrated needs to be converted to the device range accuracy specification. For example:

Provide example for temperature

### Total Accuracy of the Dynamic Temperature Measurement

The precision of dynamic temperature measurements should take into account the accuracy of the temperature sensors, transmitters, transmission means, or input to a remote readout device (such as analog input calibration to an electronic flow computer), in addition to the accuracy of the calibration reference device used.

To evaluate the measurement error of a complete temperature measurement system including the sensor, the signal transmission, display and the analog input to a flow computation device, the process of temperature calibration should include a comparison of the flow temperature measured by the sensor and the reference device located by the sensor against the temperature read and displayed by the flow computation device.

### Inspection, Verification, and Calibration Requirements

### General

All temperature measurement devices used for custody transfer shall be inspected, verified and calibrated periodically. The verification and subsequent calibration should be performed against reference standards. Just as flow meters are proved against a reference standard, temperature devices must be verified frequently as part of normal operations to assure proper performance. Necessary equipment should be installed/provided to frequently and routinely verify the operational temperature devices, which should include the measured value displayed by the local readout and remote readout.

All inspection, verification, and calibration test data including as-found values shall be recorded and preserved for inspection and audit purposes.

Annex A is a comparison of Inspection, Verification and Calibration Requirements.

NOTE	The terms “verification” and “calibration” may or may not refer to the same process in various standards.

### Traceability

Field or lab standards used for calibration shall be traceable to a national metrological institute such as NIST in that parameter, range, and scope and evidence shall be documented and provided.

### Inspection

##### General

Temperature devices shall be inspected to ensure they are in good working condition, fully functional and safe to use.

In addition, the condition of the test thermowell should be visually checked for leaks and the presence of appropriate thermal transfer media such as gels or glycol.

##### Liquid-in-Glass Thermometer

Liquid-in-glass thermometer shall be inspected upon receipt and prior to initial use. Afterwards, the thermometer shall be inspected before each use, or once per day, whichever is less frequent.

Inspection shall include visual inspection for scale markings, liquid column separation as a minimum. In addition as a check (inspection), compare two or more operational LIGTs or compare the operational LIGT to a reference thermometer at one temperature. Acceptable difference between the two should be reasonably close, depending on the quality of the devices and their accuracy. Refer to ASTM E77, Standard Test Method for Inspection and Verification of Thermometers.

A liquid-in-glass thermometer shall be handled with care and examined before each use for breakage and separation of the liquid column. Additionally, the following should be taken into consideration:

- After use in heavy oils, all parts of the thermometer assembly should be cleaned to avoid the formation of an insulating film of oil.

- A thermometer is particularly susceptible to the formation of an insulating film, so the glass thermometer should be removed frequently from the metal assembly and cleaned.

- A thermometer which has lost a substantial amount of pigment from its engraved scale should not be used in light colored products because it is difficult to read the graduation increments and numbers. If the scale graduations are not legible, replace the device.

- A thermometer with a separated liquid column should never be used. If the liquid column has been completely rejoined, the thermometer may be used, provided it is found by verification to be accurate within the tolerance.

##### Digital Thermometer

A digital thermometer shall be inspected upon receipt and prior to initial use. Afterwards, it shall be inspected to ensure they are in good working condition before each use or once per day, whichever is less frequent.

In addition, the digital thermometer may also be checked by comparing the ambient reading against a reference thermometer or against another suitable reference (such as a dry block) at one temperature. If the readings differ by more than ±0.25 °C (±0.5 °F), the digital thermometer should be recalibrated before it is used.

##### Temperature Transmitter

A newly installed temperature transmitter shall be inspected for integrity of the wiring and terminal blocks and then functionally tested. After inspection and testing, the temperature transmitter shall be calibrated. Refer to the calibration procedure in 10.5.

A temperature transmitter with a sensor shall be installed in a thermowell. Upon installation, loop check, and calibration, the transmitter is inspected to ensure it is in good operating condition and has proper sealing for weather protection and security.

##### Resistance Temperature Detector

An in-line stand-alone resistance temperature detector (RTD) sensor shall be installed in a thermowell. A newly installed RTD sensor shall be inspected for integrity of the wiring and terminal blocks and corrosion of the terminals. After inspection, the RTD output shall be calibrated against a reference device. The RTD plus line impedance if needed is offset so that the value agrees with the reference. Refer to the calibration procedure in 10.5.

Bias Adjustments: RTDs have a linearization curve and are often calibrated with a bias adjustment. If an RTD calibration requires a bias adjustment, the temperature should be verified at a second point to confirm that there is no span related calibration error.

Afterwards, periodic inspection shall include visual check for the physical condition of the sensor, cable, and local display, if any.

### Verification

##### General

For the purpose of this standard, the term “verification” refers to the process of comparing the working (operational) temperature device against a certified reference thermometer or a thermometer traceable to a national metrology institute (e.g. NIST-certified reference thermometer) to ensure accuracy or verification against a redundant temperature device.

Verification confirms whether the temperature device is operating within acceptable parameters, or requires calibration, repair or replacement.

If the tolerance levels of Annex A are exceeded, then the temperature device shall be recalibrated or replaced.

When a verification test is performed, the “as-found” reading shall be compared to the reference thermometer. Should the “as-found” reading not agree with the reference thermometer, within the required control limit, a calibration, repair or replacement of that device will be required. Whenever temperature data is transmitted to another device, the purpose of the verification process is to ensure that the measured temperature is verified at every point at which the temperature is utilized as source data for volume calculations.

Before verification of the temperature, visually inspect the field temperature device(s) to confirm good working conditions and functionality.  If the functionality of the device is questionable, issues shall be addressed prior to proceeding.

Verification (In-situ, Ex-situ)

At a minimum, verifications shall be performed at one point, as close to operating temperature as practical. The verification shall be performed in-situ using a test thermowell, thermometer and flowing fluid temperature, only if fluid is flowing, or ex-situ using a temperature bath and test thermometer near the expected operating condition.

For devices that provide a signal to a flow computation device, e.g., flow computer, the verification shall include the complete loop all the way to the flow computation device.

Consider manufacturer guidance depending on the equipment and application.

If verification fails to meet the verification tolerance requirements, the device shall be calibrated or replaced.

Verification is performed when a device is placed into service and subsequently at required intervals.

Redundancy Verification

Daily, weekly or monthly percent of reading comparisons are performed between the primary and check redundant transmitter may be used in place of periodic verification. Should the redundancy verification fail to meet its tolerance requirements, the in-situ or ex-situ verification process shall be followed.

##### Initial verification – Prior in Service

Initial verifications shall be performed when a temperature device is first placed into service. The device shall have previously been calibrated by the manufacturer, during project acceptance testing and or during commissioning. If the device does not have a valid calibration prior to being placed into service a full calibration shall be performed. Follow the manufacturer and, where available, industry standard

commissioning and calibration procedures.

For temperature devices transmitting signals to a flow computation device, e.g., flow computer (Transmitter and RTD),  verification shall include the complete loop all the way to the final device. Use the manufacturer guidance depending on what’s being used.

##### Sub sequent verification

##### Liquid-in-Glass Thermometer

A liquid-in-glass thermometer shall be verified before each use, or once per day, whichever is less frequent.

A verification shall be accomplished by placing the operational liquid-in-glass thermometer and the reference thermometer into a source of uniform and stable temperature. Using a temperature-controlled source [e.g. temperature bath or an electrical temperature calibrator (dry block)] may do this. Compare the difference between the operational liquid-in-glass thermometer and the calibration reference at near ambient temperature. If the difference exceeds the required tolerance of the glass thermometer under test and the reference specified in Annex A, the glass thermometer shall be calibrated (see 10.5) or replaced.

Prior-to-use Verification
Thermometers that are used for intermittent service, such as tank gauging, shall be inspected before each use or once per day, whichever is less frequent, to verify the thermometer is not broken and the liquid column is still intact. On a quarterly basis, the thermometer shall be verified against an NMI traceable test device of suitable resolution and accuracy at a single point, such as ambient temperature, or in a temperature-controlled source (e.g. temperature bath).

##### Digital Thermometer (PET or DCT)

A digital thermometer shall be verified monthly. The verification is accomplished by placing the operational digital thermometer and the reference thermometer into a source of uniform and stable temperature. Using a temperature- controlled source (e.g. temperature bath or an electrical temperature calibrator (metal “dry block”) may do this. Compare the difference between the operational digital thermometer and the reference thermometer at two points either:

- within the normal operating temperatures, or

- one point near ambient temperature and the other at ice point.

If the tolerance levels of Annex A are exceeded, then the temperature device shall be recalibrated or replaced.

##### Temperature Transmitter

Temperature transmitters shall be verified periodically at an interval not to exceed three months.  As a minimum verifications shall be done at one point, as close to operating temperature as practical. The verification shall be done preferably using a test thermowell, and reference thermometer under flowing conditions near normal operating conditions

. If product is not flowing accurate insitu verification is not possible and alternative verifications methods, ex-situ, are recommended.

Ex-situ verification may be performed using a bath and test thermometer. The verification is accomplished by placing the RTD and the reference thermometer into a liquid of uniform temperature, which can be a temperature bath, close to each other and allowing both to stabilize before recording the readings

Verification may be deferred, by agreement of all parties, until flow is established, however verification spacing shall not exceed six months or exceed regulatory requirements.

The verification procedure should consist of at least one comparison, preferably at the normal operating temperature, between the temperature transmitter and the reference thermometer. The verification is accomplished by placing the RTD and the reference thermometer into a liquid of uniform temperature, which can be a temperature bath or the operating pipe or vessel/container, close to each other and allowing both to stabilize before recording the readings.

Redundancy Verifications: Daily, weekly or monthly percent of reading comparisons are performed between the

primary and check redundant transmitter may be used in place of periodic verification. Should the redundancy

verification fail to meet its tolerance requirements, the verification process shown above shall be followed. (Annex I in Ch 21.1 provides an example of a Redundancy Verification Report.) Calibration requirements remain in force and apply to all redundant devices.

The verification procedure should consist of at least one comparison, preferably at the normal operating temperature, between the temperature transmitter and the reference thermometer. The verification is accomplished by placing the RTD and the reference thermometer into a liquid of uniform temperature, which can be a temperature bath or the operating pipe or vessel/container, close to each other and allowing both to stabilize before recording the readings.

The need to replace an RTD cannot be determined from a system verification as the verification involves an assessment of the entire temperature measurement system (RTD, transmitter and readout). The need to replace an RTD can only be established after the calibration of the temperature transmitter and the readout device or assessment of the RTD.

If replacement of any device of the temperature measurement system is required, then the system must be recalibrated.

If verification fails to meet the verification tolerance requirements, the device shall be calibrated or

replaced.

##### Resistance Temperature Detector

An in-line stand-alone resistance temperature detector (RTD) which is used without a temperature transmitter shall be verified periodically at an interval not to exceed three months.

Verification of the RTD (temperature sensor, without a transmitter) shall be either:

- In-situ by comparing the local readout or the remote readout (on a flow computer, electronic preset or metering control system) against a reference thermometer (such as a liquid-in-glass thermometer or digital thermometer) in a test thermowell, or

- Alternately, the RTD may be retrieved from the line and the probe (with the cabling connected) be verified against a reference thermometer (liquid-in-glass thermometer or digital thermometer) using a reference medium meeting the required uncertainty, e.g., dry block, or liquid bath, or ice point.

- If the RTD fails verification the RTD shall be replaced.

### Calibration

##### General

For the purpose of this standard, the term calibration refers to the process of verifying and then adjusting, if needed, the working temperature device output signal or reading to agree with the certified calibration reference used to within acceptable tolerance. If the device is not within acceptable tolerance, then the recalibration is to be done in accordance with the manufacturer recommendations. For a temperature device which cannot be adjusted, a correction table shall be generated from the calibration.

Depending on the temperature device type, the application, or the contractual agreement, a certificate may be required upon a successful calibration.

Reference devices used in calibration procedures for a temperature device shall be traceable to a national metrology institute, such as NIST. The requirements are typically more rigorous than those for a routine verification, for example more calibration temperature points and more stringent tolerance.

As-found and as-left readings from a calibration shall be recorded and the report may also compare them with established tolerance.

##### Tolerance Comparisons between Working and Reference Devices

An operating device is considered acceptable for custody transfer measurement if the difference between the reading of the operating device and the certified reference thermometer measurement device is within tolerance. For electronic type temperature devices, the tolerance for the comparison may range from 0.1 °C (0.2 °F) to 0.05 °C (0.1 °F).

The tolerances depend on the intended application of the working temperature device and the type of temperature device. The working temperature device needs to meet or exceed the minimum accuracy requirements in Section 9.

These tolerances may also be subject to other guidelines existing in agreements, contracts, or regulations.

##### Minimum Number of Calibration Points

Calibration of an existing or new temperature device should consist of temperature comparisons, at a minimum of three temperatures on uniform intervals over the operating range of the device. The test temperatures should be near the upper, middle and lower level of the operating range.

NOTE  Exception may warrant less calibration point(s) due to the narrow operating temperature and difficulty to pull out an inline device (such as stand-alone RTD). On the other hand, more calibration points may be needed for applications in which higher accuracy is needed over a wider range.

The calibration can be accomplished similar to the verification procedure, except that a temperature-controlled source will facilitate the three-point test. The calibration corrections must be applied to the reference thermometer readings.

##### Liquid-in-Glass Thermometer

Liquid-in-glass thermometer used in dynamic custody transfer measurement shall be calibrated annually. See API

MPMS Chapter 7.1 on calibration requirements.

##### Digital Thermometer

A digital thermometer used for dynamic custody transfer measurement shall be calibrated annually. For a digital thermometer that can’t be adjusted, a correction table should be generated from the calibration. Refer to ASTM E2593, ASTM E2877 or NIST Special Publication 250-22 on the calibration procedure of a digital thermometer.

Calibration of a digital thermometer shall be carried out in controlled temperatures.

##### Temperature Transmitter

Temperature transmitters used in dynamic custody transfer measurement shall be calibrated annually.

Each component of a temperature system (RTD, transmitter and readout) shall be verified and calibrated or replaced as required to bring the measurement system into tolerance.

The temperature sensor for an electronic temperature transmitter usually cannot be adjusted. Only the transmitter can be electrically adjusted or calibrated. A simulated input from a device such as an RTD simulator or a decade box may be used to facilitate calibration.

The result of calibration after adjustment should be verified by comparisons, within the tolerance, at a minimum of three temperatures on approximately uniform intervals over the operating range of the temperature transmitter.

The calibration procedure may vary for the specific type (analog or digital) of electronic temperature transmitter and should follow the manufacturer's recommendations.

##### Resistance Temperature Detector (RTD)

An in-line stand-alone resistance temperature detector (RTD) used in dynamic custody transfer measurement should be calibrated annually.

For the resistance temperature detector, ASTM E644 describes the test methods for insulation resistance, calibration, immersion error, pressure effects, thermal response time, vibration effect, mechanical shock, self-heating effect, stability, thermoelectric effect, humidity, thermal hysteresis, thermal shock, and end seal integrity.

Out of tolerance deviation between the working device and calibration reference may be compensated by applying an offset in the RTD readout/control unit.

If during a calibration, it has been determined that the output of the RTD needs offset adjustment, the difference between the RTD and the reference reading after the offset (i.e. post-calibration) shall agree to be within 0.1 °C if SI unit display or within 0.2 °F if U.S. customary unit display is used. A tighter calibration tolerance maybe required for applications which demand a higher accuracy.

### Certification

Upon completion of the calibration process, the individual who performed the calibration shall furnish a report stating the equipment has met or exceeded the specifications and tolerances intended to permit the use of the equipment in normal field operations or as a standard test device for other equipment.

A certificate should be issued by the company that performed the calibration. If that is the owner or operator of the facilities, concerned individuals shall be allowed to witness the calibration. The certificate may be a certified calibration report. The certification for a liquid-in-glass thermometer, a digital thermometer, or a temperature transmitter, shall have the necessary information to verify the traceability of the reference standard device used, method used, results, the individual who performed the calibration, and the expiration date of the certification.

# Annex A

(normative)

# Comparison of Inspection, Verification and Calibration Requirements

### A.1	General

This annex is intended to be a reference guide on the recommended inspection, verification, and calibration requirements for the different types of temperature devices used in dynamic temperature measurement.

##### Table A.1—Comparison of Inspection, Verification and Calibration Requirements

| Working Device | Liquid-in-Glass Thermometer | Digital Thermometer | Temperature Transmitter | Resistance Temperature Detector |
| --- | --- | --- | --- | --- |
| Process |  |  |  |  |
| Inspection | Frequency | Frequency | Frequency | Frequency |
|  | 1) Upon receipt and prior to initial use. | 1) Upon receipt and prior to initial use. | Upon installation, loop check, and calibration. | Upon installation and periodically thereafter. |
|  | 2) Daily or before each use, whichever is less frequent. | 2) Daily or before each use, whichever is less frequent. |  |  |
|  | How | How | How | How |
|  | —Physical, visual inspection —Compare two or more operational LIGTs or compare the operational LIGT to a reference thermometer at one temperature. | —Visual inspection (such as battery charged and cable integrity) and check for proper function —Two or more operational devices can be compared as a check (inspection). Inspection may include comparing the digital thermometer against a reference thermometer in liquid, or against another suitable reference (such as a dry block) at one temperature. | —Newly installed transmitter shall be inspected for proper installation and then functionally tested. —After inspection and testing, it shall be calibrated. —The transmitter shall be inspected to ensure it is in good operating condition and properly sealed for weather protection and security. | —Visual inspection for integrity of the wiring and terminal blocks for a newly installed RTD. —Afterwards, periodic inspection shall include visual check for the physical condition of the element, cable, and local display. —Compare against a reference device, if needed use offset so that RTD output agrees with the reference. |
|  | Tolerance | Tolerance | Tolerance | Tolerance |
|  | Tolerance depends on the quality and accuracy of reference and operational devices. Refer to ASTM E77, Standard, Test Method for Inspection and Verification of Thermometers. | If the readings differ by more than ±0.25 °C (0.5 °F), the digital thermometer should be recalibrated before it is used. | N/A | N/A |

| Working Device | Liquid-in-Glass Thermometer | Digital Thermometer | Temperature Transmitter | Resistance Temperature Detector |
| --- | --- | --- | --- | --- |
| Process |  |  |  |  |
| Verification | Frequency | Frequency | Frequency | Frequency |
|  | Daily or before each use, whichever less frequent. | Monthly | Every three months | Every three months |
|  | How | How | How | How |
|  | —Compare the difference between the operational liquid-in- glass thermometer and the calibration reference at near ambient temperature. | —Compare digital thermometer readout against a reference thermometer at two points; within the normal operating temperature, or one point near ambient temperature, and the other at ice point. | Compare the transmitter output or reading against the reference thermometer for at least one point, preferably at the normal operating temperature. —The verification is accomplished by placing the RTD and the reference thermometer into a liquid of uniform temperature, which can be a temperature bath or the operating pipe or vessel/container. | Verification shall be either: —In-situ by comparing the local readout or the remote readout (on a flow computer, electronic preset, or metering control system) against a reference thermometer, such as a glass thermometer or digital thermometer, in a test thermowell, or —Alternately, the RTD may be retrieved from the line and the probe (with the cabling connected) be verified against a reference thermometer (glass thermometer, digital thermometer, or dry- block). —The verification temperature may be either near ambient, or near product temperature, or at ice point. |

##### Table A.1—Comparison of Inspection, Verification and Calibration Requirements (Continued)

| Working Device | Liquid-in-Glass Thermometer | Digital Thermometer | Temperature Transmitter | Resistance Temperature Detector |
| --- | --- | --- | --- | --- |
| Process |  |  |  |  |
| Verification | Tolerance | Tolerance | Tolerance | Tolerance |
|  | If the difference exceeds the required tolerance of the glass thermometer under test and the reference specified in Annex A, the glass thermometer shall be calibrated (see 10.4.1) or replaced. | - Not to exceed the accuracy specification for the digital thermometer. - Normally not to exceed ±0.1 °C (0.2 °F). | The need to replace an RTD cannot be determined from a verification as the verification involves an assessment of the entire temperature measurement system (RTD, transmitter and readout). The need to replace an RTD can only be established after the calibration of the temperature transmitter and the readout device. If the tolerance is exceeded, then calibration of the temperature transmitter or replacement of the stand alone resistance temperature detector, or both, (e.g. RTD) is needed. If replacement of the RTD is required, then the device must be recalibrated. The temperature transmitter (including RTD and readout) shall be recalibrated if temperature from the readout and the temperature from the reference thermometer differ by more than application tolerance specified in Table 2. | If the difference exceeds the combined uncertainty of the working device and the reference thermometer, investigate the cause to determine whether it is due to wiring, line impedance compensation, or the RTD, or the display. Calibrate the RTD by determining and then applying proper offset. See calibration procedure in 10.4.5. |
| Calibration -— Calibration against reference at multiple temperatures — Certificate with test results upon successful calibration -— See “boiler plate” paragraph in the standard on calibration and tolerance. See 10.5.1 to 10.5.3. | Frequency | Frequency | Frequency | Frequency |
|  | Calibrate annually. | Calibrate annually. | Calibrate annually. | Calibrate annually. |

Table A.1—Comparison of Inspection, Verification and Calibration Requirements (Continued)

| Working Device | Liquid-in-Glass Thermometer | Digital Thermometer | Temperature Transmitter | Resistance Temperature Detector |
| --- | --- | --- | --- | --- |
| Process |  |  |  |  |
| Calibration | How | How | How | How |
|  | Liquid-in-glass thermometer used in dynamic custody transfer measurement shall be calibrated annually. See API MPMS Chapter 7.1 on calibration requirements. | Digital thermometer used for dynamic custody transfer measurement shall be calibrated annually. For a digital thermometer that can't be adjusted, a correction table should be generated from the calibration. Refer to ASTM E2593 or ASTM E2877 on the calibration procedure of a digital thermometer. | —The electronic analog or digital output of the transmitter, (e.g. 4-20 mA) is calibrated using a device such as an RTD simulator or a decade box to simulate the resistance of a RTD. —The calibration should be verified by comparisons, within the tolerance, at a minimum of three temperatures on approximately uniform intervals over the operating range of the temperature transmitter —Each component of a temperature system (RTD, transmitter and readout) shall be verified and calibrated or replaced as required to bring the measurement system into tolerance. | —ASTM E644 describes the test methods for insulation resistance, calibration, immersion error, pressure effects, thermal response time, vibration effect, mechanical shock, self- heating effect, stability, thermoelectric effect, humidity, thermal hysteresis, thermal shock, and end seal integrity. —Out of tolerance deviation between the working device and calibration reference may be compensated by applying an offset in the RTD readout/control unit. |
|  | Tolerance | Tolerance | Tolerance | Tolerance |
|  | Accuracy tolerance as specified in ASTM E1 or other applicable industry standard or as stated in Table 2. These tolerances may also be subject to other guidelines existing in agreements, contracts, or regulations. | The tolerances depend on the intended application of the working temperature device and the type of temperature device. The working temperature device needs to meet or exceed the minimum accuracy requirements in Table 2. | The tolerances depend on the intended application of the working temperature device and the type of temperature device. The working temperature device needs to meet or exceed the minimum accuracy requirements in Table 2. These tolerances may also be subject to other guidelines existing in agreements, contracts, or regulations. | The tolerances depend on the intended application of the working temperature device and the type of temperature device. The working temperature device needs to meet or exceed the minimum accuracy requirements in Table 2. These tolerances may also be subject to other guidelines existing in agreements, contracts, or regulations. |

Annex B Normative:

# Equipment Specifications

#### RTD Simulators (Decade Boxes)

An RTD simulator (decade box) shall be provided for simulating a 3-wire or 4-wire Resistance Temperature Detector (RTD) in conjunction with the calibration and troubleshooting of temperature transmitters. RTD simulators (decade boxes) used to calibrate temperature transmitters that utilize 100-Ohm (PT100) RTD’s shall meet or exceed the following specifications:

Each RTD simulator (decade box) shall be purchased with a certificate of calibration issued by an ISO/IEC 17025 accredited calibration agency. The calibration shall be traceable to an National standard and shall confirm that the RTD simulator (decade box) meets the required accuracy specification.

| 100=Ohm (PT-100) RTD Simulator (Decade Box) Specifications |  |
| --- | --- |
| Range | 10.000 to 1111.110 Ohms |
| Resolution | 0.001 Ohms |
| Accuracy | ±0.01% of Reading |
| Stability over 12 months | ±0.0035% |

#### Temperature Calibration Baths

Temperature calibration baths shall have a minimum operating range of 32 to 212 °F (0 to 100 °C). The combined uncertainty from the instability and non-uniformity in the temperature of the bath shall be equal to or less than +0.067 °F (0.037 °C) (approximately one-third of the acceptance tolerance for digital contact thermometers). The temperature calibration bath shall have sufficient depth to permit the required immersion for the reference and working temperature measurement devices as follows:

A depth of approximately 12 inches will be sufficient for all measurement devices to be verified. Each temperature calibration bath shall be suitable for use in the applicable hazardous area classification.:

| Temperature Device | Minimum Immersion |
| --- | --- |
| Reference Digital Thermometer | ¾ of the probe length |
| Reference Liquid-in-Glass Thermometer | Just below verification temperature (entire liquid column) |
| Liquid-in-Glass Thermometer | Just below verification temperature (entire liquid column) |
| Digital Thermometer | Full length of probe |
| Density Test Thermometer | Just below verification temperature (entire liquid column) |
| Sample Thermometer | ¾ of the probe length |

#### Dewar Flasks

The vessels used in the verification of thermometers shall meet the requirements for Dewar flasks with vacuum jacketed insulation as specified in DIN 12492. Dewar flasks shall be cylindrical with a minimum internal diameter of 2.75 inches (70 millimeters) and a minimum depth of 11.75 inches (300 millimeters). Dewar flasks shall be constructed from borosilicate glass meeting the requirements of ISO/DIN 3585, shall be fully silvered, and shall have an outer aluminum protective casing. The flasks shall be suitable for use with liquids having temperatures between 32 °F and 212 °F (0 °C to 100 °C). The flasks shall be furnished with a polyethylene or equivalent plastic stopper. Flasks may be furnished with or without handles.

#### Reference Liquid-in-Glass Thermometers

Reference liquid-in-glass thermometers used to verify other thermometers and instruments shall conform to the following requirements:

API Manual of Petroleum Measurement Standards (MPMS), Chapter 7, “Temperature Determination” and Chapter 7.1, “Temperature Determination-Liquid-in-Glass Thermometers”

- American Society of Testing & Materials Standard E1, “Standard Specification for ASTM Liquid-in-Glass Thermometers”

- American Society of Testing & Materials Standard E2251, “Standard Specification for Liquid-in-Glass ASTM Thermometers with Low-Hazard Precision Liquids”, or the Scientific Apparatus Makers Association

- Each reference liquid-in-glass thermometer shall be purchased with a certification that it has been calibrated in accordance with American Society of Testing & Materials Standard E77, “Standard Test Method for the Inspection and Verification of Thermometers” by an ISO/IEC 17025 accredited calibration agency. The calibration shall be traceable to an NIST standard, shall confirm that the reference liquid-in-glass thermometer meets the required accuracy specification and shall establish the correction factors to be used for verifications of field temperature measurement devices.

#### Field Digital Thermometers

Field digital thermometers used to verify field instruments shall conform to the following requirements:

- American Society of Testing & Materials Standard E2877, “Standard Guide for Digital Contact Thermometers”.

- Each reference digital thermometer shall employ a platinum resistance sensor (thermometer) (PRT) or thermistor, have a display which registers the temperature to at least 2 decimal places and have a minimum Class C accuracy, +0.10 °F (+0.05 °C).

- Each reference digital thermometer shall has a nameplate which specifies the manufacturer, model number and serial number.

- Each reference digital thermometer shall be purchased with a certification that it has undergone calibration in accordance with American Society of Testing & Materials Standard E2593, “Standard Practice for Verification of Industrial Platinum Resistance Thermometers” by an ISO/IEC 17025 accredited calibration agency. The calibration shall be traceable to a national metrology standard, shall confirm that the digital thermometer meets the required accuracy specification and shall establish any correction factors to be used for verification of other temperature measurement devices.

#### Reference Digital Thermometers

Reference digital thermometers used to verify other thermometers shall conform to the following requirements:

- American Society of Testing & Materials Standard E2877, “Standard Guide for Digital Contact Thermometers”.

- Each reference digital thermometer employs a platinum resistance sensor (thermometer) (PRT) or thermistor, has a display which registers the temperature to at least 3 decimal places and has a minimum Class B accuracy, +0.04 °F (+0.02 °C).

- Each reference digital thermometer has a nameplate which specifies the manufacturer, model number and serial number.

- Each reference digital thermometer shall be purchased with a certification that it has undergone calibration in accordance with American Society of Testing & Materials Standard E2593, “Standard Practice for Verification of Industrial Platinum Resistance Thermometers” by an ISO/IEC 17025 accredited calibration agency. The calibration shall be traceable to a national metrology standard, shall confirm that the digital thermometer meets the required accuracy specification and shall establish any correction factors to be used for verification of other temperature measurement devices.

#### Reference Digital Multimeters, Loop Calibrators and (Multifunction) Process Calibrators

A precision digital multimeter, loop calibrator or process calibrator shall be used as a reference device for calibrating the analog outputs from field transmitters and the analog inputs to flow computers. Instruments will meet the following minimum requirements:

- Each reference digital multimeter, loop calibrator and process calibrator shall have an accuracy of +0.0150% of reading plus 2 times the last significant digit (LSD) or better. Each reference digital multimeter, loop calibrator and process calibrator shall have a resolution of 0.001 mA when used to measure (or source) analog instrument signals (e.g., 4-3/4 digits or 40000 counts).

- Each reference digital multimeter, loop calibrator or process calibrator shall be suitable for the area classification where it will be used.

- Each reference digital multimeter, loop calibrator or process calibrator shall be purchased with a certificate of calibration issued by an ISO/IEC 17025 accredited calibration agency. The calibration shall be traceable to an NIST standard and shall confirm that the digital multimeter, loop calibrator or process calibrator meets the required accuracy specification.

Annex C Normative:

# RTD Tolerance vs. Test Temperature Table

Sensor specifications require use of ASTM E1137 Class A platinum RTD sensors. A tolerance chart for a Class A 100 Ω 0.00385 curve platinum RTD and associated equation is provided below for convenience.

Grade A = ± [0.13 + 0.0017|t|] °C

# Tolerance Chart for ASTM E1137 Class A

| ASTM E 1137 |  |  |  |
| --- | --- | --- | --- |
| 100 Ω 0.00385 PRT Class A |  |  |  |
| Temperature |  | Tolerance |  |
| °C | °F | °C | °F |
| 0 | 32 | 0.13 | 0.23 |
| 5 | 41 | 0.14 | 0.25 |
| 10 | 50 | 0.15 | 0.26 |
| 15 | 59 | 0.16 | 0.28 |
| 20 | 68 | 0.16 | 0.30 |
| 25 | 77 | 0.17 | 0.31 |
| 30 | 86 | 0.18 | 0.33 |
| 35 | 95 | 0.19 | 0.34 |
| 40 | 104 | 0.20 | 0.36 |
| 45 | 113 | 0.21 | 0.37 |
| 50 | 122 | 0.22 | 0.39 |
| 55 | 131 | 0.22 | 0.40 |
| 60 | 140 | 0.23 | 0.42 |
| 65 | 149 | 0.24 | 0.43 |
| 70 | 158 | 0.25 | 0.45 |
| 75 | 167 | 0.26 | 0.46 |
| 80 | 176 | 0.27 | 0.48 |
| 85 | 185 | 0.27 | 0.49 |
| 90 | 194 | 0.28 | 0.51 |
| 95 | 203 | 0.29 | 0.52 |
| 100 | 212 | 0.30 | 0.54 |

| Description |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Platinum RTD Resistance vs. Temperature Tables |  |  |  |  |  |  |  |  |  |  |
| Resistance (Ohms) vs. Temperature (°C)* 100-Ohm Platinum (PT100) RTD with Alpha (α) of 0.00385 Reference Resistance of 32 °C |  |  |  |  |  |  |  |  |  |  |
| Temp (°C) | +0 | +1 | +2 | +3 | +4 | +5 | +6 | +7 | +8 | +9 |
| 0 | 100.00 | 100.39 | 100.78 | 101.17 | 101.56 | 101.95 | 102.34 | 102.73 | 103.12 | 103.51 |
| 10 | 103.90 | 104.29 | 104.68 | 105.07 | 105.46 | 105.85 | 106.24 | 106.63 | 107.02 | 107.41 |
| 20 | 107.79 | 108.18 | 108.57 | 108.96 | 109.35 | 109.74 | 110.12 | 110.51 | 110.90 | 111.29 |
| 30 | 111.67 | 112.06 | 112.45 | 112.84 | 113.22 | 113.61 | 114.00 | 114.38 | 114.77 | 115.16 |
| 40 | 115.54 | 115.93 | 116.31 | 116.70 | 117.09 | 117.47 | 117.86 | 118.24 | 118.63 | 119.01 |
| 50 | 119.40 | 119.78 | 120.17 | 120.55 | 120.94 | 121.32 | 121.71 | 122.09 | 122.48 | 122.86 |
| 60 | 123.25 | 123.63 | 124.01 | 124.40 | 124.78 | 125.16 | 125.55 | 125.93 | 126.31 | 126.70 |
| 70 | 127.08 | 127.46 | 127.84 | 128.23 | 128.61 | 128.99 | 129.37 | 129.76 | 130.14 | 130.52 |
| 80 | 130.90 | 131.28 | 131.66 | 132.05 | 132.43 | 132.81 | 133.19 | 133.57 | 133.95 | 134.33 |
| 90 | 134.71 | 135.09 | 135.47 | 135.85 | 136.23 | 136.61 | 136.99 | 137.37 | 137.75 | 138.13 |
| 100 | 138.51 |  |  |  |  |  |  |  |  |  |

| Description |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Platinum RTD Resistance vs. Temperature Tables |  |  |  |  |  |  |  |  |  |  |
| Resistance (Ohms) vs. Temperature (°F)* 100-Ohm Platinum (PT100) RTD with Alpha (α) of 0.00385 Reference Resistance of 32 °F |  |  |  |  |  |  |  |  |  |  |
| Temp (°F) | +0 | +1 | +2 | +3 | +4 | +5 | +6 | +7 | +8 | +9 |
| 0 | 93.03 | 93.25 | 93.47 | 93.69 | 93.91 | 94.12 | 94.34 | 94.56 | 94.78 | 95.00 |
| 10 | 95.21 | 95.43 | 95.65 | 95.87 | 96.09 | 96.30 | 96.52 | 96.74 | 96.96 | 97.17 |
| 20 | 97.39 | 97.61 | 97.83 | 98.04 | 98.26 | 98.48 | 98.70 | 98.91 | 99.13 | 99.35 |
| 30 | 99.57 | 99.78 | 100.00 | 100.22 | 100.43 | 100.65 | 100.87 | 101.09 | 101.30 | 101.52 |
| 40 | 101.74 | 101.95 | 102.17 | 102.39 | 102.60 | 102.82 | 103.04 | 103.25 | 103.47 | 103.69 |
| 50 | 115.06 | 104.12 | 104.34 | 104.55 | 104.77 | 104.98 | 105.20 | 105.42 | 105.63 | 105.85 |
| 60 | 106.07 | 106.28 | 106.50 | 106.71 | 106.93 | 107.15 | 107.36 | 107.58 | 107.79 | 108.01 |
| 70 | 108.23 | 108.44 | 108.66 | 108.87 | 109.09 | 109.30 | 109.52 | 109.73 | 109.95 | 110.17 |
| 80 | 110.38 | 110.60 | 110.81 | 111.03 | 111.24 | 111.46 | 111.67 | 111.89 | 112.10 | 112.32 |
| 90 | 112.53 | 112.75 | 112.96 | 113.18 | 113.39 | 113.61 | 113.82 | 114.04 | 114.25 | 114.47 |
| 100 | 114.68 | 114.90 | 115.11 | 115.33 | 115.54 | 115.76 | 115.97 | 116.18 | 116.40 | 116.61 |
| 110 | 116.83 | 117.04 | 117.26 | 117.47 | 117.69 | 117.90 | 118.11 | 118.33 | 118.54 | 118.76 |
| 120 | 118.97 | 119.18 | 119.40 | 119.61 | 119.83 | 120.04 | 120.25 | 120.47 | 120.68 | 120.89 |
| 130 | 121.11 | 121.32 | 121.54 | 121.75 | 121.96 | 122.18 | 122.39 | 122.60 | 122.82 | 123.03 |
| 140 | 123.24 | 123.46 | 123.67 | 123.88 | 124.10 | 124.31 | 124.52 | 124.74 | 124.95 | 125.16 |
| 150 | 125.37 | 125.59 | 125.80 | 126.01 | 126.23 | 126.44 | 126.65 | 126.86 | 127.08 | 127.29 |
| 160 | 127.50 | 127.71 | 127.93 | 128.14 | 128.35 | 128.56 | 128.78 | 128.99 | 129.20 | 129.41 |
| 170 | 129.63 | 129.84 | 130.05 | 130.26 | 130.47 | 130.69 | 130.90 | 131.11 | 131.32 | 131.53 |
| 180 | 131.75 | 131.96 | 132.17 | 132.38 | 132.59 | 132.81 | 133.02 | 133.23 | 133.44 | 133.65 |
| 190 | 133.86 | 134.08 | 134.29 | 134.50 | 134.71 | 134.92 | 135.13 | 135.34 | 135.55 | 135.77 |
| 200 | 135.98 |  |  |  |  |  |  |  |  |  |
| Platinum (PT1000) RTD R0 1000.00 Α 0.0039083 Β -5.775E-07 C -4.183E-12 Celsius (°C) VS Resistance Ω |  |  |  |  |  |  |  |  |  |  |
| (°C) | +0 | +1 | +2 | +3 | +4 | +5 | +6 | +7 | +8 | +9 |
| -50 | 803.06 | 807.03 | 811.00 | 814.97 | 818.94 | 822.90 | 826.87 | 830.83 | 834.79 | 838.75 |
| -40 | 842.71 | 846.66 | 850.62 | 854.57 | 858.53 | 862.48 | 866.43 | 870.38 | 874.32 | 878.27 |
| -30 | 882.22 | 886.16 | 890.10 | 894.04 | 897.98 | 901.92 | 905.86 | 909.80 | 913.73 | 917.67 |
| -20 | 921.60 | 925.53 | 929.46 | 933.39 | 937.32 | 941.24 | 945.17 | 949.09 | 953.02 | 956.94 |
| -10 | 960.86 | 964.78 | 968.70 | 972.61 | 976.53 | 980.44 | 984.36 | 988.27 | 992.18 | 996.09 |
| 0 | 1000.00 | 1003.91 | 1007.81 | 1011.72 | 1015.62 | 1019.53 | 1023.43 | 1027.33 | 1031.23 | 1035.13 |
| 10 | 1039.03 | 1042.92 | 1046.82 | 1050.71 | 1054.60 | 1058.50 | 1062.39 | 1066.28 | 1070.16 | 1074.05 |
| 20 | 1077.94 | 1081.82 | 1085.71 | 1089.59 | 1093.47 | 1097.35 | 1101.23 | 1105.11 | 1108.99 | 1112.86 |
| 30 | 1116.74 | 1120.61 | 1124.48 | 1128.36 | 1132.23 | 1136.09 | 1139.96 | 1143.83 | 1147.70 | 1151.56 |
| 40 | 1155.42 | 1159.29 | 1163.15 | 1167.01 | 1170.87 | 1174.73 | 1178.58 | 1182.44 | 1186.29 | 1190.15 |
| 50 | 1194.00 | 1197.85 | 1201.70 | 1205.55 | 1209.39 | 1213.24 | 1217.09 | 1220.93 | 1224.77 | 1228.61 |
| 60 | 1232.46 | 1236.29 | 1240.13 | 1243.97 | 1247.81 | 1251.64 | 1255.47 | 1259.31 | 1263.14 | 1266.97 |
| 70 | 1270.79 | 1274.62 | 1278.45 | 1282.27 | 1286.10 | 1289.92 | 1293.74 | 1297.56 | 1301.38 | 1305.19 |
| 80 | 1309.01 | 1312.83 | 1316.64 | 1320.45 | 1324.26 | 1328.07 | 1331.88 | 1335.69 | 1339.49 | 1343.30 |
| 90 | 1347.10 | 1350.90 | 1354.70 | 1358.50 | 1362.30 | 1366.09 | 1369.89 | 1373.68 | 1377.47 | 1381.27 |
| 100 | 1385.06 | 1388.84 | 1392.63 | 1396.41 | 1400.20 | 1403.98 | 1407.76 | 1411.54 | 1415.32 | 1419.09 |
| 110 | 1422.87 | 1426.64 | 1430.41 | 1434.19 | 1437.95 | 1441.72 | 1445.49 | 1449.25 | 1453.01 | 1456.78 |
| 120 | 1460.54 | 1464.29 | 1468.05 | 1471.80 | 1475.56 | 1479.31 | 1483.06 | 1486.81 | 1490.56 | 1494.30 |
| 130 | 1498.04 | 1501.79 | 1505.53 | 1509.26 | 1513.00 | 1516.74 | 1520.47 | 1524.20 | 1527.93 | 1531.66 |
| 140 | 1535.38 | 1539.11 | 1542.83 | 1546.55 | 1550.27 | 1553.99 | 1557.70 | 1561.42 | 1565.13 | 1568.84 |
| 150 | 1572.55 | 1576.25 | 1579.96 | 1583.66 | 1587.36 | 1591.06 | 1594.75 | 1598.45 | 1602.14 | 1605.83 |

| Platinum (PT1000) RTD R0 1000.00 Α 0.0039083 Β -5.775E-07 C -4.183E-12 Fahrenheit (°F) VS Resistance Ω |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| (°F) | +0 | +1 | +2 | +3 | +4 | +5 | +6 | +7 | +8 | +9 |
| -40 | 842.71 | 844.90 | 847.10 | 849.30 | 851.50 | 853.69 | 855.89 | 858.09 | 860.28 | 862.48 |
| -30 | 864.67 | 866.87 | 869.06 | 871.25 | 873.45 | 875.64 | 877.83 | 880.02 | 882.22 | 884.41 |
| -20 | 886.60 | 888.79 | 890.98 | 893.17 | 895.36 | 897.55 | 899.74 | 901.92 | 904.11 | 906.30 |
| -10 | 908.49 | 910.67 | 912.86 | 915.04 | 917.23 | 919.41 | 921.60 | 923.78 | 925.97 | 928.15 |
| 0 | 930.33 | 932.52 | 934.70 | 936.88 | 939.06 | 941.24 | 943.42 | 945.61 | 947.79 | 949.97 |
| 10 | 952.14 | 954.32 | 956.50 | 958.68 | 960.86 | 963.04 | 965.21 | 967.39 | 969.57 | 971.74 |
| 20 | 973.92 | 976.09 | 978.27 | 980.44 | 982.62 | 984.79 | 986.97 | 989.14 | 991.31 | 993.48 |
| 30 | 995.66 | 997.83 | 1000.00 | 1002.17 | 1004.34 | 1006.51 | 1008.68 | 1010.85 | 1013.02 | 1015.19 |
| 40 | 1017.36 | 1019.53 | 1021.70 | 1023.86 | 1026.03 | 1028.20 | 1030.36 | 1032.53 | 1034.70 | 1036.86 |
| 50 | 1039.03 | 1041.19 | 1043.35 | 1045.52 | 1047.68 | 1049.85 | 1052.01 | 1054.17 | 1056.33 | 1058.50 |
| 60 | 1060.66 | 1062.82 | 1064.98 | 1067.14 | 1069.30 | 1071.46 | 1073.62 | 1075.78 | 1077.94 | 1080.10 |
| 70 | 1082.25 | 1084.41 | 1086.57 | 1088.73 | 1090.88 | 1093.04 | 1095.20 | 1097.35 | 1099.51 | 1101.66 |
| 80 | 1103.82 | 1105.97 | 1108.12 | 1110.28 | 1112.43 | 1114.58 | 1116.74 | 1118.89 | 1121.04 | 1123.19 |
| 90 | 1125.34 | 1127.49 | 1129.65 | 1131.80 | 1133.95 | 1136.09 | 1138.24 | 1140.39 | 1142.54 | 1144.69 |
| 100 | 1146.84 | 1148.98 | 1151.13 | 1153.28 | 1155.42 | 1157.57 | 1159.72 | 1161.86 | 1164.01 | 1166.15 |
| 110 | 1168.29 | 1170.44 | 1172.58 | 1174.73 | 1176.87 | 1179.01 | 1181.15 | 1183.29 | 1185.44 | 1187.58 |
| 120 | 1189.72 | 1191.86 | 1194.00 | 1196.14 | 1198.28 | 1200.42 | 1202.55 | 1204.69 | 1206.83 | 1208.97 |
| 130 | 1211.10 | 1213.24 | 1215.38 | 1217.51 | 1219.65 | 1221.78 | 1223.92 | 1226.05 | 1228.19 | 1230.32 |
| 140 | 1232.46 | 1234.59 | 1236.72 | 1238.85 | 1240.99 | 1243.12 | 1245.25 | 1247.38 | 1249.51 | 1251.64 |
| 150 | 1253.77 | 1255.90 | 1258.03 | 1260.16 | 1262.28 | 1264.41 | 1266.54 | 1268.67 | 1270.79 | 1272.92 |
| 160 | 1275.05 | 1277.17 | 1279.30 | 1281.42 | 1283.55 | 1285.67 | 1287.79 | 1289.92 | 1292.04 | 1294.16 |
| 170 | 1296.29 | 1298.41 | 1300.53 | 1302.65 | 1304.77 | 1306.89 | 1309.01 | 1311.13 | 1313.25 | 1315.37 |
| 180 | 1317.49 | 1319.60 | 1321.72 | 1323.84 | 1325.96 | 1328.07 | 1330.19 | 1332.30 | 1334.42 | 1336.53 |
| 190 | 1338.65 | 1340.76 | 1342.87 | 1344.99 | 1347.10 | 1349.21 | 1351.32 | 1353.44 | 1355.55 | 1357.66 |
| 200 | 1359.77 | 1361.88 | 1363.99 | 1366.09 | 1368.20 | 1370.31 | 1372.42 | 1374.53 | 1376.63 | 1378.74 |
| 210 | 1380.84 | 1382.95 | 1385.06 | 1387.16 | 1389.26 | 1391.37 | 1393.47 | 1395.57 | 1397.68 | 1399.78 |
| 220 | 1401.88 | 1403.98 | 1406.08 | 1408.18 | 1410.28 | 1412.38 | 1414.48 | 1416.58 | 1418.68 | 1420.77 |
| 230 | 1422.87 | 1424.97 | 1427.06 | 1429.16 | 1431.25 | 1433.35 | 1435.44 | 1437.54 | 1439.63 | 1441.72 |
| 240 | 1443.81 | 1445.91 | 1448.00 | 1450.09 | 1452.18 | 1454.27 | 1456.36 | 1458.45 | 1460.54 | 1462.62 |
| 250 | 1464.71 | 1466.80 | 1468.88 | 1470.97 | 1473.06 | 1475.14 | 1477.23 | 1479.31 | 1481.39 | 1483.48 |
| 260 | 1485.56 | 1487.64 | 1489.72 | 1491.80 | 1493.88 | 1495.96 | 1498.04 | 1500.12 | 1502.20 | 1504.28 |
| 270 | 1506.36 | 1508.43 | 1510.51 | 1512.59 | 1514.66 | 1516.74 | 1518.81 | 1520.88 | 1522.96 | 1525.03 |
| 280 | 1527.10 | 1529.17 | 1531.24 | 1533.31 | 1535.38 | 1537.45 | 1539.52 | 1541.59 | 1543.66 | 1545.72 |
| 290 | 1547.79 | 1549.86 | 1551.92 | 1553.99 | 1556.05 | 1558.12 | 1560.18 | 1562.24 | 1564.30 | 1566.36 |
| 300 | 1568.43 | 1570.49 | 1572.55 | 1574.60 | 1576.66 | 1578.72 | 1580.78 | 1582.83 | 1584.89 | 1586.95 |
| 310 | 1589.00 | 1591.06 | 1593.11 | 1595.16 | 1597.21 | 1599.27 | 1601.32 | 1603.37 | 1605.42 | 1607.47 |
| 320 | 1609.52 | 1611.56 | 1613.61 | 1615.66 | 1617.70 | 1619.75 | 1621.79 | 1623.84 | 1625.88 | 1627.93 |

Annex D Informative:

# Lumped Capacitance Model (First-Order Approximation)

Estimating the response time of a thermometer in a thermowell (often called "thermal response time" or "time constant") involves understanding how quickly the thermometer reacts to changes in fluid temperature. The response time depends on:

- Thermowell material and thickness

- Thermometer type

- Fluid velocity and properties

- Heat transfer coefficient

### 1. Lumped Capacitance Model (First-Order Approximation)

For most practical cases, the system can be modeled as a first-order thermal system:

τ=ρVCphA\tau = \frac{\rho V C_p}{h A}τ=hAρVCp​​

Where:

- τ\tauτ = time constant (seconds)

- ρ\rhoρ = density of thermowell material (kg/m³)

- VVV = volume of thermowell (m³)

- CpC_pCp​ = specific heat of thermowell material (J/kg·K)

- hhh = convective heat transfer coefficient (W/m²·K)

- AAA = surface area exposed to fluid (m²)

The response time is typically defined as the time to reach 63.2% of the final temperature after a step change.

Annex E Informative:

# Response Time Estimation

Practical Steps to Estimate Response Time

- Determine Thermowell Properties:

- Material (e.g., stainless steel)

- Dimensions (length, diameter, wall thickness)

- Estimate Heat Transfer Coefficient (hhh):

- Use empirical correlations for forced convection (e.g., Dittus-Boelter for turbulent flow):

Nu=0.023Re0.8Pr0.4Nu = 0.023 Re^{0.8} Pr^{0.4}Nu=0.023Re0.8Pr0.4

h=Nu⋅kDh = \frac{Nu \cdot k}{D}h=DNu⋅k​

Where NuNuNu is Nusselt number, ReReRe is Reynolds number, PrPrPr is Prandtl number, kkk is fluid thermal conductivity, DDD is diameter.

- Calculate Time Constant (τ\tauτ):

- Plug values into the lumped capacitance formula.

- Estimate Response Time:

- For a step change, time to reach 99% of final value is about 5τ5\tau5τ.

3. Example Calculation

Suppose:

- Stainless steel thermowell, ρ=8000\rho = 8000ρ=8000 kg/m³, Cp=500C_p = 500Cp​=500 J/kg·K

- Volume V=1×10−5V = 1 \times 10^{-5}V=1×10−5 m³

- Surface area A=2×10−3A = 2 \times 10^{-3}A=2×10−3 m²

- h=500h = 500h=500 W/m²·K

τ=8000×1×10−5×500500×2×10−3=401=40 seconds\tau = \frac{8000 \times 1 \times 10^{-5} \times 500}{500 \times 2 \times 10^{-3}} = \frac{40}{1} = 40 \text{ seconds}τ=500×2×10−38000×1×10−5×500​=140​=40 seconds

Time to reach 99% response: 5×40=2005 \times 40 = 2005×40=200 seconds.

# Bibliography

- API MPMS Chapter 2, Upright Cylindrical Tanks

- API MPMS Chapter 11.1, Physical Properties Data—Temperature and Pressure Volume Correction Factors for Generalized Crude Oils, Refined Products, and Lubricating Oils

- API MPMS Chapter 15, Guidelines for Use of the International System of Units (SI) in the Petroleum and Allied Industries

American Society of Mechanical Engineers, Two Park Avenue, New York, NY 10016-5990

ASTM International, 100 Barr Harbor Drive, West Conshohocken, Pennsylvania 19428, www.astm.org

International Association of Classification Societies, IACS Limited, 6th Floor, 36 Broadway, London SW1H 0BH, United Kingdom, www.iacs.org.uk

National Fire Protection Association, 1 Batterymarch Park, Quincy, Massachusetts 02269, www.nfpa.org

National Institute of Standards and Technology, 100 Bureau Drive, Gaithersburg, Maryland 20899, www.nist.gov