Manual of Petroleum Measurement Standards Chapter 10.4

Determination of Water and/or Sediment in Crude Oil by the Centrifuge Method (Field Procedure)

FIFTH EDITION, AUGUST 2020

REAFFIRMED, NOVEMBER 2025

Special Notes

API publications necessarily address problems of a general nature. With respect to particular circumstances, local, state, and federal laws and regulations should be reviewed. The use of API publications is voluntary. In some cases, third parties or authorities having jurisdiction may choose to incorporate API standards by reference and may mandate compliance.

Neither API nor any of API’s employees, subcontractors, consultants, committees, or other assignees make any warranty or representation, either express or implied, with respect to the accuracy, completeness, or usefulness of the information contained herein, or assume any liability or responsibility for any use, or the results of such use, of any information or process disclosed in this publication. Neither API nor any of API’s employees, subcontractors, consultants, or other assignees represent that use of this publication would not infringe upon privately owned rights.

API publications may be used by anyone desiring to do so. Every effort has been made by the Institute to ensure the accuracy and reliability of the data contained in them; however, the Institute makes no representation, warranty, or guarantee in connection with this publication and hereby expressly disclaims any liability or responsibility for loss or damage resulting from its use or for the violation of any authorities having jurisdiction with which this publication may conflict.

API publications are published to facilitate the broad availability of proven, sound engineering and operating practices. These publications are not intended to obviate the need for applying sound engineering judgment regarding when and where these publications should be used. The formulation and publication of API publications is not intended in any way to inhibit anyone from using any other practices.

Any manufacturer marking equipment or materials in conformance with the marking requirements of an API standard is solely responsible for complying with all the applicable requirements of that standard. API does not represent, warrant, or guarantee that such products do in fact conform to the applicable API standard.

Classified areas may vary depending on the location, conditions, equipment, and substances involved in any given

situation. Users of this Standard should consult with the appropriate authorities having jurisdiction.

Users of this Standard should not rely exclusively on the information contained in this document. Sound business,

scientific, engineering, and safety judgment should be used in employing the information contained herein.

API is not undertaking to meet the duties of employers, manufacturers, or suppliers to warn and properly train and equip their employees, and others exposed, concerning health and safety risks and precautions, nor undertaking their obligations to comply with authorities having jurisdiction.

Information concerning safety and health risks and proper precautions with respect to particular materials and conditions should be obtained from the employer, the manufacturer or supplier of that material, or the material safety data sheet.

All rights reserved. No part of this work may be reproduced, translated, stored in a retrieval system, or transmitted by any means, electronic, mechanical, photocopying, recording, or otherwise, without prior written permission from the publisher. Contact the Publisher, API Publishing Services, 200 Massachusetts Avenue, NW, Suite 1100, Washington, DC 20001-5571.

Copyright © 2020 American Petroleum Institute

ii

Foreword

Nothing contained in any API publication is to be construed as granting any right, by implication or otherwise, for the manufacture, sale, or use of any method, apparatus, or product covered by letters patent. Neither should anything contained in the publication be construed as insuring anyone against liability for infringement of letters patent.

The verbal forms used to express the provisions in this document are as follows.

Shall: As used in a standard, “shall” denotes a minimum requirement to conform to the standard.

Should: As used in a standard, “should” denotes a recommendation or that which is advised but not required to conform to the standard.

May: As used in a standard, “may” denotes a course of action permissible within the limits of a standard. Can: As used in a standard, “can” denotes a statement of possibility or capability.

This document was produced under API standardization procedures that ensure appropriate notification and participation in the developmental process and is designated as an API standard. Questions concerning the interpretation of the content of this publication or comments and questions concerning the procedures under which this publication was developed should be directed in writing to the Director of Standards, American Petroleum Institute, 200 Massachusetts Avenue, NW, Suite 1100, Washington, DC 20001. Requests for permission to reproduce or translate all or any part of the material published herein should also be addressed to the director.

Generally, API standards are reviewed and revised, reaffirmed, or withdrawn at least every five years. A one-time extension of up to two years may be added to this review cycle. Status of the publication can be ascertained from the API Standards Department, telephone (202) 682-8000. A catalog of API publications and materials is published annually by API, 200 Massachusetts Avenue, NW, Suite 1100, Washington, DC 20001.

Suggested revisions are invited and should be submitted to the Standards Department, API, 200 Massachusetts Avenue, NW, Suite 1100, Washington, DC 20001, standards@api.org.

iii

Contents

Page

- Scope	1

- Normative References	1

- Terms and Definitions	2

- Significance and Use	2

- Test Solvents	3

  - General	3

  - Use of Solvents	3

- Demulsifier	3

- Apparatus	3

  - Centrifuge	3

  - Centrifuge Tubes	4

  - Preheater	5

  - Sample Thermometer	5

- Sampling	5

- Procedure	5

  - General	5

  - Determination of Sediment and Water	8

  - Determination of Sediment Only	11

- Calculation and Reporting of Results	13

  - General	13

  - 100 mL Tubes	13

  - 200-part Tubes	13

- Precision	14

- Bias	14

Annex A (informative) Precautionary Information	15

Annex B (informative) Demulsifiers and Water Saturation of Solvents	18

Annex C (normative) Specifications for Centrifuge Tubes	19

Annex D (informative) Abbreviated Procedure & Calculations—Determination of Sediment & Water	23

Annex E (informative) Abbreviated Procedure & Calculations—Determination of Sediment Only	24

Bibliography	25

v

Contents

Page

Figures

- Reading a 100 mL Centrifuge Tube	7

- Reading a 200-part Centrifuge Tube	8

  - 8 in. (203 mm) Cone-shaped Centrifuge Tube	20

  - 6 in. (167 mm) Cone-shaped Centrifuge Tube	21

### Tables

- Rotation Speeds Necessary to Produce a Relative Centrifugal Force of 500 g for Centrifuges

of Various Diameters of Swing	4

- Procedure for Reading a 100 mL Cone-shaped Tube	7

- Procedure for Reading a 200-part Cone-shaped Tube	8

- Allowable Difference between Results from Two 100 mL Tubes	11

- Allowable Difference between Results from Two 200-part Tubes	11

- Calculation and Reporting of Test Results	13

  - Minimum Graduation Requirements and Maximum Calibration Tolerances for 8 in. (203 mm)

Cone-shaped Tubes	21

  - Minimum Graduation Requirements and Maximum Calibration Tolerances for 6 in. (167 mm)

Cone-shaped Tubes	22

vi

Determination of Water and/or Sediment in Crude Oil by the Centrifuge Method (Field Procedure)

## Scope

      - This section describes the field centrifuge method for determining both sediment and water or sediment only in crude oil. This method may not always produce the most accurate results, but it is considered the most practical method for field determination of sediment and water. This method may also be used for field determination of sediment. When a higher degree of accuracy is required, the laboratory procedure described in API MPMS Ch. 10.3, Standard Test Method for Water and Sediment in Crude Oil by the Centrifuge Method (Laboratory Procedure) (ASTM D4007); API MPMS Chapter 10.2, Standard Test Method for Water in Crude Oil by Distillation (ASTM D4006); or API MPMS Chapter 10.9, Standard Test Method for Water in Crude Oils by Coulometric Karl Fischer Titration (ASTM D4928); and the procedure described in API MPMS Chapter 10.1, Standard Test Method for Sediment in Crude Oils and Fuel Oils by the Extraction Method (ASTM D473) or API MPMS Chapter 10.8, Standard Test Method for Sediment in Crude Oil by Membrane Filtration (ASTM D4807) should be used.

      - The requirements of this standard may involve hazardous materials, operations, and equipment. This standard does not purport to address all of the safety problems associated with its use. The user of this standard has the responsibility for establishing appropriate safety and health practices and determining the applicability of regulatory limitations prior to use.

## Normative References

The following documents are referred to in the text in such a way that some or all of their content constitutes requirements of this document. For dated references, only the edition cited applies. For undated references, the latest edition of the referenced document (including any addenda) applies.

API MPMS Chapter 8.1, Manual Sampling of Petroleum and Petroleum Products

API MPMS Chapter 8.2, Automatic Sampling of Petroleum and Petroleum Products

API MPMS Chapter 8.3, Standard Practice for Mixing and Handling of Liquid Samples of Petroleum and Petroleum Products (includes Errata 1 dated March 1996)

API MPMS Chapter 10.1, Standard Test Method for Sediment in Crude Oils and Fuel Oils by the Extraction Method

API MPMS Chapter 10.2, Standard Test Method for Water in Crude Oil by Distillation

API MPMS Chapter 10.3, Standard Test Method for Water and Sediment in Crude Oil by the Centrifuge Method (Laboratory Procedure)

API MPMS Chapter 10.8, Standard Test Method for Sediment in Crude Oil by Membrane Filtration

API MPMS Chapter 10.9, Standard Test Method for Water in Crude Oils by Coulometric Karl Fischer Titration ASTM D473,1 Standard Test Method for Sediment in Crude Oils and Fuel Oils by the Extraction Method ASTM D3699, Standard Specification for Kerosine

ASTM D4006, Standard Test Method for Water in Crude Oil by Distillation

1  ASTM International, 100 Barr Harbor Drive, West Conshohocken, Pennsylvania 19428, www.astm.org.

1

- API Manual of Petroleum Measurement Standards 10.4

ASTM D4007, Standard Test Method for Water and Sediment in Crude Oil by the Centrifuge Method (Laboratory Procedure)

ASTM D4057, Standard Practice for Manual Sampling of Petroleum and Petroleum Products ASTM D4177, Standard Practice for Automatic Sampling of Petroleum and Petroleum Products ASTM D4807, Standard Test Method for Sediment in Crude Oil by Membrane Filtration

ASTM D4928, Standard Test Method for Water in Crude Oils by Coulometric Karl Fischer Titration

ASTM D5854, Standard Practice for Mixing and Handling of Liquid Samples of Petroleum and Petroleum Products (includes Errata 1 dated March 1996)

ASTM E542, Standard Practice for Calibration of Laboratory Volumetric Apparatus

## Terms and Definitions

For the purposes of this document, the following definitions apply.

3.1

sediment

A material that normally exists in crude oils as extremely fine, well-dispersed solids that originate in the reservoir from which the crude came, or in drilling fluids used to drill the wells, which may take the form of sand, clay, shale, or rock particles. Other sediments such as scale can be picked up from tubing, pipe, tanks, and other production and transportation equipment. Sediment, as defined in this standard, shall have a density greater than water and will not float at test conditions.

3.2

demulsifier-stock solution

A mixture of demulsifier and solvent in a fixed ratio (e.g. 25 % to 75 %) that is added to the test solvent to form a

solution to cause a clear and distinct separation between the oil and water interface.

3.3

demulsifier-solvent solution

A mixture of any one of the approved solvents and a demulsifier, or a demulsifier-stock solution that is used as a

diluent to help the separation of sediment and water from the sample being tested.

3.4

solvent

Any one of the approved liquids capable of dissolving the hydrocarbon fraction of the crude oil to form a solution. A solvent is added to reduce the viscosity of the crude oil and promote separation of the crude oil hydrocarbons from the sediment and water in the sample being tested.

## Significance and Use

A determination of water and/or sediment content is required to accurately determine the net volumes of crude oil involved in sales, taxation, exchanges, inventories, and custody transfers. An excessive amount of water and/or sediment in crude oil is significant because it can cause corrosion of equipment and problems in processing and transporting, and may violate federal, state, or municipal regulations.

For determination of both sediment and water, known volumes of crude oil, demulsifier, and solvent (water-saturated if required) are placed in a centrifuge tube and heated to 60 °C ± 3 °C (140 °F ± 5 °F). After centrifuging, the volumes of sediment and water at the bottom of the tube are read. For some crude oils, temperatures of 71 °C (160 °F) or higher may be required to completely melt wax crystals in crude oils. Temperatures shall be maintained above the cloud point of the crude oil being tested.

Manual of Petroleum Measurement Standards Chapter 10.4	3

For determination of sediment only, known volumes of crude oil, solvent, demulsifiers, and water are placed in a centrifuge tube and heated to 71 °C ± 3 °C (160 °F ± 5 °F). After centrifuging, the volume of the sediment at the bottom of the tube is read.

## Test Solvents

### General

Kerosine, Stoddard solvent, toluene, and xylene are satisfactory for use in field testing. Solvents and demulsifiers are mixed with samples in the centrifuge tubes to reduce the viscosity and to allow sediment and water to separate and fall to the tip of the tube during the testing procedure. Also, solvents are used to help dissolve paraffin and asphaltenes to prevent them from being centrifuged to the tip of the tube and included in the test results. Different types of crude oil may require the use of different solvents to obtain accurate results. For example, kerosine and Stoddard solvent will not dissolve asphaltenes.

Testing should be performed to determine the best solvent to use on a particular crude type.

Annex A contains general information regarding the physical characteristics of these solvents, as well as fire, safety, and health considerations relating to their use. Kerosine shall meet the specifications of ASTM D3699.

Information concerning physical characteristics, safety and health risks, and proper precautions with respect to particular materials and conditions should be obtained from the employer, the manufacturer, or the supplier of that material, or the material safety data sheet (MSDS).

### Use of Solvents

Toluene and xylene shall be saturated with water at the test temperature used and maintained at that temperature until used. A procedure for the saturation of solvents is given in Annex B. The water-saturated solvent shall be free from suspended water at the time of use. Toluene or xylene is recommended for sediment and water determinations involving asphaltenic crude oils.

Stoddard solvent and kerosine do not have to be saturated with water, since the solubility of water in these

solvents is not significant at the test temperature.

## Demulsifier

Unless the results of testing indicate that a demulsifier is not required for a particular crude type, demulsifiers shall be used to promote the separation of water from the sample and make it easier to read the water/oil interface. Demulsifiers should be mixed according to the manufacturer’s recommendations and should not add to the volume of sediment and water determined.

The demulsifier shall always be used in the form of a demulsifier stock solution or be premixed with the solvent to be used in the test. The ratio of demulsifier to solvent should be recorded to be able to duplicate results if necessary. Testing should be done to determine which demulsifier provides the best results with a specific crude type. See Annex A and B for further information about demulsifiers.

## Apparatus

### Centrifuge

A centrifuge shall be capable of spinning two or more centrifuge tubes at a speed that can be controlled to give a minimum relative centrifugal force of 500 g at the tip of the tubes. The rotation speed necessary to produce a relative centrifugal force of 500 g for various diameters of swing can be determined from Table 1 or from the following equation:

_

rpm = α √

_rcf

d

(1)

4	API Manual of Petroleum Measurement Standards 10.4

where

rpm	is the rotation speed in revolutions per minute;

rcf	is the relative centrifugal force in multiples of gravity;

d	is the diameter of swing, in millimeters or inches, measured between the tips of opposite tubes when the tubes are in their fully extended rotating position;

α	is 1335 rpm for d in millimeters, and 265 rpm for d in inches.

Table 1—Rotation Speeds Necessary to Produce a Relative Centrifugal Force of 500 g for Centrifuges of Various Diameters of Swing

| Diameter of swing “d ” a |  | Rotation Speed (Revolutions per Minute) |
| --- | --- | --- |
| Millimeters | Inches |  |
| 305 | 12 | 1710 |
| 330 | 13 | 1640 |
| 356 | 14 | 1580 |
| 381 | 15 | 1530 |
| 406 | 16 | 1480 |
| 432 | 17 | 1440 |
| 457 | 18 | 1400 |
| 483 | 19 | 1360 |
| 508 | 20 | 1325 |
| 533 | 21 | 1290 |
| 559 | 22 | 1260 |
| 584 | 23 | 1240 |
| 610 | 24 | 1210 |
| a Measured between the tips of opposite tubes when the tubes are in their fully extended rotating position. |  |  |

The revolving head, trunnion rings, and trunnion cups, including the cushions, shall be constructed to withstand the maximum centrifugal force capable of being delivered by the power source. The trunnion cups and cushions shall firmly support the tubes when the centrifuge is in motion. The centrifuge shall be enclosed by a metal shield or case strong enough to contain flying debris in the event a tube breaks or the centrifuge malfunctions.

The centrifuge machine shall be heated and be capable of maintaining the sample within 8 °C (15 °F) of the test temperature. For example, the minimum acceptable temperature after centrifuging is 52 °C (125 °F) at a test temperature of 60 °C (140 °F) or 63 °C (145 °F) at a test temperature of 71 °C (160 °F).

### Centrifuge Tubes

Centrifuge tubes shall be cone-shaped and 8 in. (203 mm) or 6 in. (167 mm) in length. Tubes shall conform to the dimensions given in Annex C and shall be made of thoroughly annealed glass. The tube graduation marks’ accuracy shall be volumetrically verified or gravimetrically certified before field use of the tube, in accordance with ASTM E542, using National Institute of Standards and Technology (NIST) or other internationally accepted traceable equipment.

The 6 in. tubes are more commonly used in field applications due to the smaller size and weight of the centrifuge required. The 8 in. tubes are easier to read due to the increased spacing of the graduations. In either case, the tubes can be 100 mL or 200-part tubes. The computation methods for the total percent of sediment and water are different based on the type of tube selected. For 100 mL tubes, the sediment and water volume readings from

Manual of Petroleum Measurement Standards Chapter 10.4	5

each tube are added. For 200-part tubes, the sediment and water volume readings from each tube are averaged. See Figure 1 for an example of a 100 mL tube and Table 2 for how to read a 100 mL tube. See Figure 2 for an example of a 200-part tube and Table 3 for how to read a 200-part tube.

### Preheater

The preheater shall be either a metal block or a liquid bath of sufficient depth to permit immersion of the centrifuge tube in the vertical position to the 100 mL (200-part) mark and capable of heating the sample to 71 °C ± 3 °C (160 °F ± 5 °F).

### Sample Thermometer

A sample thermometer shall have graduations at intervals of 1 °C (2 °F) or less and shall be accurate to ± 0.5 °C (± 1 °F). A bimetal, dial-type thermometer is normally used.

## Sampling

Sampling is defined as all steps required to obtain a representative quantity of the contents of any pipeline, tank, or other system and the placement of it, with care, in an appropriate analytical device, such as a centrifuge tube, to maintain its compositional integrity prior to testing.

Samples obtained and handled as specified in API MPMS Ch. 8.1 (ASTM D4057), Ch. 8.2 (ASTM D4177), and Ch. 8.3 (ASTM D5854) shall be used for this method.

## Procedure

### General

The centrifuge sediment and water test consists of the following basic activities:

    - dispensing of crude oil aliquots from the same previously mixed crude oil sample into two centrifuge tubes;

    - addition of solvent and demulsifier to each tube;

    - heating the contents of the respective tubes to a designated test temperature;

    - centrifuging the contents of the tubes;

    - reading of the results from the respective tubes;

    - calculation and reporting of the final test results (Section 10).

For the test to be valid, the following conditions shall be met:

- The temperatures of the contents of the respective tubes shall be controlled within specified limits.

- The oil/water interface for each tube shall be clear and distinct and no distinguishable layer shall be present above the interface.

- The results for each tube following two successive centrifuging operations shall agree.

- The results from the respective tubes shall agree within a specified tolerance.

When the crude oil aliquots are introduced into the tubes prior to the solvent, care should be taken to ensure that precisely 50 mL (100 parts) of crude oil are dispensed into each tube. In this case, the volume of the solvent (and

6	API Manual of Petroleum Measurement Standards 10.4

demulsifier) is less critical as it only acts as a diluent and any inaccuracy in the quantity of solvent added will not affect the test results.

In cases where the crude oil is very viscous, the solvent may be dispensed into each tube prior to the crude oil to enhance mixing. In such cases, care should be taken to ensure each centrifuge tube is filled exactly to the 50 mL (100-part) mark with solvent and then exactly to the 100 mL (200-part) mark with crude oil.

Due to the opaqueness of the liquid, a centrifuge tube shall be read at the top of the meniscus when establishing the quantity of crude oil introduced into the tube. It is recognized that reading the top of the meniscus will result in a bias error in the test result that may or may not be significant, depending upon the water and/or sediment content. The bias error in the observed sediment and water reading (mL or parts) may be approximated from an estimate of the volume (mL) associated with the depth of the meniscus at the 50 mL (100-part) mark as follows:

Error in Reading (mL) = Reading (mL) * ⟨1 – {50 / [50 – Meniscus Volume (mL) ] } ⟩

Error in Reading (parts) = Reading (parts) * (1 – ⟨100 / {100 – [2 * Meniscus Volume (mL) ] } ⟩ )

Given a meniscus volume of 0.6 mL at the 50 mL (100-part) mark of a 6 in. centrifuge tube, the bias error in the sediment and water reading may be approximated as follows:

Error in Reading (mL) = (–0.01215) * Reading (mL) Error in Reading (parts) = (–0.01215) * Reading (parts)

In the event that the amount of a crude oil aliquot or solvent dispensed into a tube exceeds the required quantity, the crude oil aliquot and/or solvent shall be discarded, the tube shall be cleaned, and the procedure shall be reinitiated.

Each user should provide a means for identifying the tubes for purposes of making the required comparisons. Use of centrifuge tubes with serial numbers or centrifuge machines with unique identifiers on the respective tube compartments are two ways to meet this requirement.

The use of stoppers in each procedure is required to maintain the representativeness of the crude oil aliquots.

Manual of Petroleum Measurement Standards Chapter 10.4	7

Figure 1—Reading a 100 mL Centrifuge Tube

Table 2—Procedure for Reading a 100 mL Cone-shaped Tube

| Volume (mL) | Read to Nearest (mL) |
| --- | --- |
| 0.0 to 0.1 | 0.025 |
| 0.1 to 1.0 | 0.050 |
| 1.0 to 3.0 | 0.100 |
| 3.0 to 5.0 | 0.500 |
| >5.0 | Interpolate to Nearest Half Increment |

8	API Manual of Petroleum Measurement Standards 10.4

Figure 2—Reading a 200-part Centrifuge Tube

Table 3—Procedure for Reading a 200-part Cone-shaped Tube

| Volume (%) | Read to Nearest (%) |
| --- | --- |
| 0.0 to 0.2 | 0.05 |
| 0.2 to 2.0 | 0.10 |
| 2.0 to 6.0 | 0.20 |
| 6.0 to 10 | 1.0 |
| >10 | Interpolate to Nearest Half Increment |

### Determination of Sediment and Water

NOTE	A summary of the requirements for determining sediment and water is provided in Annex D.

- If toluene or xylene is the solvent, water-saturate the toluene or xylene at the test temperature. See Annex B or API MPMS Chapter 10.3, Annex A.1.

NOTE The degree to which toluene or xylene has an affinity for water is dependent upon the temperature. To ensure that it does not absorb water from the crude oil under test and thereby adversely affect the test results, it is critical that the toluene or xylene solvent be saturated with water within 3 °C (5 °F) of the test temperature when dispensed into the centrifuge tubes.

Manual of Petroleum Measurement Standards Chapter 10.4	9

- Fill each of two centrifuge tubes to the 50 mL (100-part) mark with a crude oil aliquot drawn directly from the sampling device or the container in which the sample was mixed. Read the top of the meniscus when measuring the amount of sample added to each tube.

- Fill each centrifuge tube to the 100 mL (200-part) mark with solvent only or a premix of solvent and demulsifier, as applicable. Read the top of the meniscus when determining the total quantity of solvent and crude oil in each tube. If a premix of solvent and demulsifier is used, take care not to add any free demulsifier. If toluene or xylene is the solvent, take care not to add any free water from the water-saturated solvent.

NOTE Introduction of free demulsifier from a premix of solvent and demulsifier, and/or free water associated with the water-saturated toluene or xylene will adversely affect the test results.

- If demulsifier has not been premixed with the solvent and testing or experience indicates that it is required, add to each tube, a quantity of demulsifier-stock solution that has previously been determined to be satisfactory for the crude oil under test. Typically, 2 drops to 4 drops of demulsifier-stock solution is appropriate.

- Tightly stopper each centrifuge tube. Invert each tube a minimum of 10 times to ensure the oil and solvent are uniformly mixed.

CAUTION—Each tube should always be inverted at a position below and away from the face and eyes so

that contact with the mixture will be avoided if the stopper becomes loose.

- Loosely secure the stopper on each tube to prevent pressure buildup during heating, and immerse each tube to the 100 mL (200-part) mark in a sample heater or water bath.

- Heat the contents of the tubes to a test temperature of 60 °C ± 3 °C (140 °F ± 5 °F). Use a sample thermometer to verify the temperature of the contents in each tube.

NOTE 1	Testing or experience may indicate that temperatures of 71 °C ± 3 °C (160 °F ± 5 °F) or more may be

necessary to ensure any paraffin is completely liquefied and promote water separation.

NOTE 2	Insufficient heating time may adversely affect the test results when using water-saturated aromatic solvents. Additional heating time may be required to ensure that any water that separates from the water-saturated solvent is fully reabsorbed by the solvent at the test temperature prior to centrifuging.

- Tighten the stopper in each tube. Invert each tube a minimum of 10 times to ensure the oil and solvent are uniformly mixed.

- Place the tubes in opposite trunnion cups in the centrifuge and spin the tubes for at least 5 minutes at a minimum relative centrifugal force of 500 g. See Equation 1 and Table 1.

NOTE Testing or experience may indicate that the spin time needs to be longer than 5 minutes, depending upon the physical properties of the crude oil being tested.

- Immediately after the centrifuge comes to rest, verify the temperature of the contents of each tube is within 8 °C (15 °F) of the test temperature with a sample thermometer, taking care not to disturb the oil/water interface. If the temperature of the contents of either tube differs from the test temperature by more than 8 °C (15 °F), adjust the temperature setting of the centrifuge and repeat the procedure from Step 6.

NOTE The test is invalid if the temperature is outside the required temperature range.

- Verify the oil/water interface is clear, distinct and readable, and no identifiable layer (e.g. emulsion, congealed paraffin, and/or asphaltenes) is present above the interface in each tube. If the interface is unclear or there is an additional layer above the interface, accomplish one or more of the following actions:

  - Without agitation, repeat the procedure from Step 6 using the same test temperature.

  - Without agitation, repeat the procedure from Step 6 using a test temperature of 71 °C ± 3 °C (160 °F ± 5 °F) if not previously accomplished.

10	API Manual of Petroleum Measurement Standards 10.4

  - Add additional demulsifier-stock solution and repeat the procedure from Step 5.

  - Reinitiate the test from Step 1 for different crude oil aliquots from the same sample.

  - Determine the sediment and water content using another API/ASTM test method(s).

NOTE The test is invalid if the oil/water interface is not clear and distinct or a distinguishable layer is present above the interface in either tube.

- While holding the tube in an upright (vertical) position, read and record the total volume of sediment and water at the bottom of each tube in accordance with the requirements specified in Figure 1 and Table 2 for 100 mL tubes, or Figure 2 and Table 3 for 200-part tubes.

- Without agitation, loosen the stopper on each tube to prevent pressure buildup during heating, and immerse each tube to the 100 mL (200-part) mark in the sample heater or water bath.

- Heat the contents of the tubes to a test temperature of 60 °C ± 3 °C (140 °F ± 5 °F). Use a sample thermometer to verify the temperature of the contents of each tube.

NOTE Testing or experience may indicate that temperatures of 71 °C ± 3 °C (160 °C ± 5 °F) or more may be necessary

to ensure any paraffin is completely liquefied and promote water separation.

- Without agitation, tighten the stopper in each tube.

- Place the two tubes in opposite trunnion cups in the centrifuge and spin the tubes for at least 5 minutes at a minimum relative centrifugal force of 500 g. See Equation 1 and Table 1.

- Immediately after the centrifuge comes to rest, verify the temperature of the contents in each tube is within 8 °C (15 °F) of the test temperature with a sample thermometer, taking care not to disturb the oil/water interface. If the temperature of the contents of either tube differs from the test temperature by more than 8 °C (15 °F), adjust the temperature setting of the centrifuge and repeat the procedure from Step 13.

NOTE The test is invalid if the temperature is outside the required temperature range.

- While holding the tube in an upright (vertical) position, read and record the total volume of sediment and water at the bottom of each tube in accordance with the requirements specified in Figure 1 and Table 2 for 100 mL tubes, or Figure 2 and Table 3 for 200-part tubes.

- Compare the last two readings for each tube. If the combined volume of sediment and water in each tube is not the same as previously determined, repeat the procedure from Step 13 until two consecutive readings agree for each tube.

- Record the final volume of sediment and water in each tube.

- Compare the test results from the two centrifuge tubes. If the difference between the results from the respective tubes is greater than the limit specified in Table 4 for 100 mL tubes or Table 5 for 200-part tubes, repeat the procedure from Step 1. Otherwise, proceed to Section 10 for final calculation and reporting of the test results.

NOTE The test is invalid if the two readings differ by more than the applicable limit from Table 4 or Table 5.

Manual of Petroleum Measurement Standards Chapter 10.4	11

Table 4—Allowable Difference between Results from Two

100 mL Tubes

| Result from Tube with Highest Reading (mL) | Acceptable Difference Between Tubes (mL) |
| --- | --- |
| 0.000 to 0.500 | 0.050 |
| Above 0.500 to 2.000 | 0.100 |

Table 5—Allowable Difference between Results from Two 200-part Tubes

| Result from Tube with Highest Reading (Parts) | Acceptable Difference Between Tubes (Parts) |
| --- | --- |
| 0.000 to 1.000 | 0.100 |
| Above 1.000 to 4.000 | 0.200 |

### Determination of Sediment Only

NOTE 1	A summary of the requirements for determining sediment only is provided in Annex E.

NOTE 2	Because the objective of this test is to determine the sediment content only, water saturation of toluene or xylene is not required. Water absorption into or dropout from the toluene or xylene has no bearing on the test results.

- Fill each of two centrifuge tubes to the 50 mL (100-part) mark with a crude oil aliquot drawn directly from the sampling device or the container in which the sample was mixed. Read the top of the meniscus when measuring the amount of sample added to each tube.

- Fill each centrifuge tube to the 100 mL (200-part) mark with solvent only or a premix of solvent and demulsifier, as applicable. Read the top of the meniscus when determining the total quantity of solvent and crude oil in each tube. If a premix of solvent and demulsifier is used, take care not to add any free demulsifier.

NOTE Introduction of free demulsifier from a premix of solvent and demulsifier will adversely affect the test results.

- If demulsifier has not been premixed with the solvent and testing or experience indicates that it is required, add to each tube, a quantity of demulsifier-stock solution that has previously been determined to be satisfactory for the crude oil under test. Typically, 2 drops to 4 drops of demulsifier-stock solution is appropriate.

- Add 3 drops to 4 drops (~0.15 mL to 0.20 mL) of demineralized or distilled water to each tube.

- Tightly stopper each centrifuge tube. lnvert each tube a minimum of 10 times to ensure the oil and solvent are uniformly mixed.

CAUTION—Each tube should always be inverted at a position below and away from the face and eyes so

that contact with the mixture will be avoided if the stopper becomes loose.

- Loosely secure the stopper on each tube to prevent pressure buildup during heating, and immerse each tube to the 100 mL mark in a sample heater or water bath.

- Heat the contents of the tubes to a test temperature of 71 °C ± 3 °C (160 °F ± 5 °F). Use a sample thermometer to verify the temperature of the contents in each tube.

- Tighten the stopper in each tube. Invert each tube a minimum of 10 times to ensure the oil and solvent are uniformly mixed.

- Place the tubes in opposite trunnion cups in the centrifuge, and spin the tubes for at least 5 minutes at a minimum relative centrifugal force of 500 g. See Equation 1 and Table 1.

12	API Manual of Petroleum Measurement Standards 10.4

NOTE Testing or experience may indicate that the spin time needs to be longer than 5 minutes, depending upon the physical properties of the crude oil being tested.

- Immediately after the centrifuge comes to rest, verify the temperature of the contents in each tube is within 8 °C (15 °F) of the test temperature with a sample thermometer, taking care not to disturb the oil/water interface. If the temperature of the contents of either tube differs from the test temperature by more than 8 °C (15 °F), adjust the temperature setting of the centrifuge and repeat the procedure from Step 6.

NOTE The test is invalid if the temperature is outside the required temperature range.

- Verify the oil/water interface is clear, distinct, and readable, and no identifiable layer (e.g. emulsion, congealed paraffin, and/or asphaltenes) is present above the interface in each tube. If the interface is unclear or there is an additional layer above the interface, accomplish one or more of the following actions:

  - Without agitation, repeat the procedure from Step 6.

  - Add additional demulsifier-stock solution and repeat the procedure from Step 5.

  - Reinitiate the test from Step 1 for different crude oil aliquots from the same sample.

  - Determine the sediment content using another API/ASTM test method(s).

NOTE The test is invalid if the oil/water interface is not clear and distinct or a distinguishable layer is present above the interface in either tube.

- While holding the tube in an upright (vertical) position, read and record the sediment volume at the bottom of each tube in accordance with the requirements specified in Figure 1 and Table 2 for 100 mL tubes, or Figure 2 and Table 3 for 200-part tubes.

NOTE If the sediment volume in a tube is less than 0.0125 mL (0.025 parts), the volume of sediment for the tube should be recorded as 0.000 mL (0.000 parts).

- Without agitation, loosen the stopper on each tube to prevent pressure buildup during heating and immerse each tube to the 100 mL (200-part) mark in the sample heater or water bath.

- Heat the contents of the tubes to a test temperature of 71 °C ± 3 °C (160 °F ± 5 °F). Use a sample thermometer to verify the temperature of the contents of each tube.

- Without agitation, tighten the stopper in each tube.

- Place the two tubes in opposite trunnion cups in the centrifuge and spin the tubes for at least 5 minutes at a minimum relative centrifugal force of 500 g. See Equation 1 and Table 1.

- Immediately after the centrifuge comes to rest, verify the temperature of the contents in each tube is within 8 °C (15 °F) of the test temperature with a sample thermometer, taking care not to disturb the oil/water interface. If the temperature of the contents of either tube differs from the test temperature by more than 8 °C (15 °F), adjust the temperature setting of the centrifuge and repeat the procedure from Step 13.

NOTE The test is invalid if the temperature is outside the required temperature range.

- While holding the tube in an upright (vertical) position, read and record the volume of sediment at the bottom of each tube in accordance with the requirements specified in Figure 1 and Table 2 for 100 mL tubes, or Figure 2 and Table 3 for 200-part tubes.

NOTE If the sediment volume in a tube is less than 0.0125 mL (0.025 parts), the volume of sediment for the tube should be recorded as 0.000 mL (0.000 parts).

- Compare the last two readings for each tube. If the volume of sediment in each tube is not the same as previously determined, repeat the procedure from Step 13 until two consecutive readings agree for each tube.

- Record the final volume of sediment in each tube.

Manual of Petroleum Measurement Standards Chapter 10.4	13

- Compare the test results from the two centrifuge tubes. If the difference between the results from the respective tubes is greater than the limit specified in Table 4 for 100 mL tubes or Table 5 for 200-part tubes, repeat the procedure from Step 1. Otherwise, proceed to Section 10 for final calculation and reporting of the test results.

NOTE The test is invalid if the two readings differ by more than the applicable limit from Table 4 or Table 5.

## Calculation and Reporting of Results

### General

The methodology for calculating the test result is dependent upon the type of tubes used.

### 100 mL Tubes

For 100 mL tubes, add the final readings from the two tubes and report the resultant sum as the percentage by volume of sediment and water or sediment only to three decimal points. Table 6 summarizes the calculation and reporting requirements for combinations of tube readings up to 0.150 mL.

### 200-part Tubes

For 200-part tubes, average the final readings from the two tubes and report the result as the percentage by volume of sediment and water or sediment only to three decimal points. Table 6 summarizes the calculation and reporting requirements for combinations of tube readings up to 0.300 parts.

Table 6—Calculation and Reporting of Test Results

| 100 mL Tubes Volume of Water and/or Sediment (mL) |  | 200-part Tubes Volume of Water and/or Sediment (Parts) |  | Reported Total Water and/or Sediment |
| --- | --- | --- | --- | --- |
| Tube 1 | Tube 2 | Tube 1 | Tube 2 | (%) |
| No visible water and sediment | No visible water and sediment | No visible water and sediment | No visible water and sediment | 0.000 |
| No visible water and sediment | 0.025 | No visible water and sediment | 0.050 | 0.025 |
| No visible water and sediment | 0.050 | No visible water and sediment | 0.100 | 0.050 |
| 0.025 | 0.025 | 0.050 | 0.050 | 0.050 |
| 0.025 | 0.050 | 0.050 | 0.100 | 0.075 |
| 0.025 | 0.075 | 0.050 | 0.150 | 0.100 |
| 0.050 | 0.050 | 0.100 | 0.100 | 0.100 |
| 0.050 | 0.075 | 0.100 | 0.150 | 0.125 |
| 0.050 | 0.100 | 0.100 | 0.200 | 0.150 |
| 0.075 | 0.075 | 0.150 | 0.150 | 0.150 |
| 0.075 | 0.100 | 0.150 | 0.200 | 0.175 |
| 0.100 | 0.100 | 0.200 | 0.200 | 0.200 |
| 0.100 | 0.150 | 0.200 | 0.300 | 0.250 |
| 0.150 | 0.150 | 0.300 | 0.300 | 0.300 |

14	API Manual of Petroleum Measurement Standards 10.4

## Precision

No precision for the test method described herein has been established. This is a field method. Experience has shown that the determination of both sediment and water using this test method has sometimes been as precise as other accepted test methods.

## Bias

Since there is no accepted reference material suitable for determining the bias for the procedure described in these sections for measuring water and/or sediment in crude oil by the centrifuge method (field procedure), no statement about bias is being made.

Annex A

(informative)

Precautionary Information

  - Physical Characteristics and Fire Considerations

Information regarding particular materials and conditions should be obtained from the employer, the manufacturer, or supplier of that material, or the MSDS.

    - Kerosine (ASTM D3699)

The typical characteristics of kerosine are a distillation range of 205 °C to 300 °C (401 °F to 572 °F), a minimum

flash point of 38 °C (100 °F), and a maximum freezing point of –30 °C (–22 °F).

CAUTION—Kerosine is combustible and should be kept away from heat, sparks, or open flame. Kerosine containers should be kept closed when they are not in use. Fires should be extinguished with foam, dry chemical, or carbon dioxide. Spills should be soaked up with clay, diatomaceous earth, or a similar material.

    - Stoddard Solvent

The typical characteristics of Stoddard solvent are a distillation range of 149 °C to 208 °C (300 °F to 407 °F), a

minimum flash point of 38 °C (100 °F), and an aromatics plus olefins content of less than 20 percent by volume.

CAUTION—Stoddard solvent is combustible and should be kept away from heat, sparks, or open flame. Stoddard solvent containers should be kept closed when they are not in use. Fires should be extinguished with foam, dry chemical, or carbon dioxide. Spills should be soaked up with clay, diatomaceous earth, or a similar material.

    - Toluene

The typical characteristics of toluene are a minimum flash point of 4 °C (40 °F), an ignition temperature of 480 °C (896 °F), a boiling point of 111 °C (231 °F), an upper flammability limit of 7.1 percent by volume, a lower flammability limit of 1.2 percent by volume, a molecular weight of 92, an American Public Health Association (APHA) color of ten (per ASTM D1209), a boiling range (initial to dry point) of 2.0 °C (3.6 °F) (recorded boiling point of 110.6 °C [231.1 °F]), and 0.001 % residue after evaporation. Toluene passes the American Chemical Society (ACS) test for substances darkened by H2SO4.

CAUTION—Toluene is flammable and should be kept away from heat, sparks, or open flame. Toluene containers should be kept closed when they are not in use. Fires should be extinguished with water spray, foam, dry chemical, or carbon dioxide. Spills should be soaked up with clay, diatomaceous earth, or a similar material.

Xylene

The typical characteristics of xylene are a minimum flash point of 27 °C (81 °F), an ignition temperature of 527 °C (982 °F), a boiling point of 139 °C (282 °F), an upper flammability limit of 7.0 percent by volume, a lower flammability limit of 1.1 percent by volume, a molecular weight of 106, an APHA color of not more than ten

15

16	API Manual of Petroleum Measurement Standards 10.4

(per ASTM D1209), a boiling range of 137 °C to 144 °C (279 °F to 291 °F), and 0.002 percent residue after evaporation. Xylene passes the ACS test for substances darkened by H2SO4.

CAUTION—Xylene is flammable and should be kept away from heat, sparks, or open flame. Xylene containers should be kept closed when they are not in use. Fires should be extinguished with water spray, foam, dry chemical, or carbon dioxide. Spills should be soaked up with clay, diatomaceous earth, or a similar material.

  - Safety and Health Considerations

    - General Considerations

Potential health effects that result from chemical exposure depend on the chemical concentration and the length

of exposure. Everyone should minimize exposure to chemicals. The following general precautions are suggested.

- Minimize skin and eye contact and breathing of vapors.

- Keep chemicals away from the mouth, since they can be harmful or fatal if swallowed or aspirated.

- Keep chemical containers closed when they are not in use.

- Keep work areas well-ventilated and as clean as possible.

- Promptly clean up spills in accordance with pertinent safety, health, and environmental regulations.

- Observe established exposure limits and use proper protective clothing and equipment.

NOTE  For permissible exposure limits (PELs), consult the most recent edition of the Occupational Safety and Health Standards, Title 29 Code of Federal Regulations (CFR) Part 1910, Subpart Z, “Toxic and Hazardous Substances.” [2] Also consult the most recent edition of the ACGIH booklet, Threshold Limit Values and Biological Exposure.[3] See the Bibliography for complete reference information.

    - Kerosine

Health effects can result from exposure to kerosine through contact with the skin and eyes, breathing of vapors, swallowing, or aspiration. Kerosine can exhibit local irritant properties that may be manifested by dermatitis of the skin; stinging of the eyes, nose, or throat; or respiratory irritation. Acute exposure to kerosine may result in adverse systemic effects, including effects on the central nervous, respiratory, or gastrointestinal systems. Symptoms of systemic effects may include vomiting, diarrhea, and in severe cases, drowsiness and central nervous system depression progressing to coma and death.

There may also be long-term (chronic) health effects of varying severity from exposure to kerosine.

Information concerning safety and health risks and proper precautions with respect to particular materials and conditions should be obtained from the employer, the manufacturer, or supplier of that material, or the MSDS.

    - Stoddard Solvent

Health effects can result from exposure to Stoddard solvent through contact with the skin and eyes, breathing of vapors, swallowing, or aspiration. Stoddard solvent exhibits local irritant properties that may be manifested by dermatitis of the skin; stinging of the eyes, nose, or throat; or respiratory irritation. Acute exposure to Stoddard solvent may result in adverse systemic effects including effects on the central nervous, respiratory, or gastrointestinal systems. Symptoms of systemic effects may include headache, fatigue, nausea, vomiting,

Manual of Petroleum Measurement Standards Chapter 10.4	17

cough, and gastrointestinal or respiratory difficulties. There may also be long-term (chronic) health effects of varying severity from exposure to Stoddard solvent. Information concerning safety and health risks and proper precautions with respect to particular materials and conditions should be obtained from the employer, the manufacturer, or supplier of that material, or the MSDS.

    - Toluene

Health effects can result from exposure to toluene through contact with the skin and eyes, breathing of vapors, swallowing, or aspiration. Toluene exhibits local irritant properties that may be manifested by dermatitis of the skin; stinging of the eyes, nose, or throat; or respiratory irritation. Acute exposure to toluene may result in adverse systemic effects including effects on the central nervous, cardiovascular, respiratory, or gastrointestinal systems.

Symptoms of acute systemic effects may include headache, fatigue, disturbed vision, dizziness, confusion, numbness of the hands and feet, heart palpitations, anorexia, and gastrointestinal or respiratory difficulties. Overexposure may lead to coma and death.

There may also be long-term (chronic) health effects of varying severity from exposure to toluene.

Information concerning safety and health risks and proper precautions with respect to particular materials and conditions should be obtained from the employer, the manufacturer, or supplier of that material, or the MSDS.

    - Xylene

Health effects can result from exposure to xylene through contact with the skin and eyes, breathing of vapors, swallowing, or aspiration. Xylene exhibits local irritant properties that may be manifested by dermatitis of the skin; stinging of the eyes, nose, or throat; or respiratory irritation. Acute exposure to xylene may result in adverse systemic effects including effects on the central nervous, cardiovascular, respiratory, or gastrointestinal systems. Symptoms of acute systemic effects may include headache, fatigue, disturbed vision, dizziness, confusion, numbness of the hands and feet, heart palpitations, anorexia, and gastrointestinal or respiratory difficulties. Overexposure may lead to coma and death.

There may also be long-term (chronic) health effects of varying severity from exposure to xylene.

Information concerning safety and health risks and proper precautions with respect to particular materials and conditions should be obtained from the employer, the manufacturer, or supplier of that material, or the MSDS.

    - Demulsifiers

Various chemical demulsifiers may be used in the procedures described in this manual.

Information concerning safety and health risks and proper precautions with respect to particular materials and conditions should be obtained from the employer, the manufacturer, or supplier of that material, or the MSDS.

Annex B

(informative)

Demulsifiers and Water Saturation of Solvents

  - Water Saturation of Solvents

    - Fill each of two centrifuge tubes to the 2 mL (4-part) mark with demineralized or distilled water and then to the 100 mL (200-part) mark with the solvent to be saturated.

    - Stopper the tubes and shake vigorously for 30 seconds, holding the tubes below eye level, to disperse the water in the solvent. Loosen the stoppers.

    - Place the tubes containing the water/solvent mixture into a sample preheater. Heat the tubes to 60 °C (140 °F) or other test temperature and maintain this temperature for a minimum of 30 minutes.

    - Inspect the water/solvent mixture for suspended water droplets before use. If any suspended water is visible, the tubes shall be centrifuged at a temperature of 60 °C (140 °F) or other test temperature for 5 minutes at a speed sufficient to give a relative centrifugal force of 500 g at the tube tip.

    - Use the top 50 mL (100-parts) of the mixture from each tube for test purposes. Take particular care not to pour any of the free water in the tip of each tube into the sample.

  - Demulsifiers

    - Although a good commercial crude oil demulsifier will work effectively with a wide range of crude oils, there are some crudes for which one demulsifier is more effective than another. Oil-soluble demulsifiers, which are generally used in this procedure, normally consist of an aromatic base (usually xylene or toluene) plus a proprietary active ingredient. Water-soluble demulsifiers normally consist of an alcohol base plus a proprietary active ingredient. For specific information regarding demulsifiers, consult Annex A and the MSDS for the particular demulsifier in use.

    - Excessive amounts of demulsifier will cause errors. Demulsifiers do not, in general, contain water, but they do have a limited solubility in the solvent-sample mix and, if added in excessive quantities, can show up after centrifuging as a separate immiscible component at the bottom of the tube. To prevent this problem, the demulsifier should always be used in the form of a demulsifier-stock solution or should be premixed with the solvent to be used in the test. The ratio of demulsifier to solvent should be recorded to be able to duplicate results using the solution if necessary. In either case, the quantity of demulsifier to be added to the solvent should be based on tests for the particular demulsifier-solvent solution combination. The demulsifier manufacturer’s instructions should be followed when a demulsifier-solvent solution is prepared.

    - Demulsifiers and demulsifier-solvent solutions should always be stored in accordance with the manufacturer’s recommendations. Each container used to store a demulsifier or demulsifier-solvent solution should be dated and the contents should be discarded when they reach the manufacturer’s recommended shelf life.

18

Annex C

(normative)

Specifications for Centrifuge Tubes

Centrifuge tubes shall be cone-shaped and 8 in. (Figure C.1) or 6 in. (Figure C.2) in length and conform to the dimensions shown. The tubes shall be made of thoroughly annealed glass.

A 100 mL tube shall conform to the dimensions shown in Figure C.1. A 200-part tube shall conform to the dimensions shown in Figure C.2, with the marking for each division multiplied by 2 (for example, 25 mL = 50 parts). The mouth of each tube shall be constricted for closure with a stopper. Graduations for the 8 in. (203 mm) and 6 in. (167 mm) tubes shall be in accordance with the requirements of Table C.1 and Table C.2, respectively. The scale errors for a centrifuge tube shall not exceed the tolerances specified in these tables. The graduation requirements and scale error tolerances shown in Table C.1 and Table C.2 apply to calibrations made by reading the bottom of the shaded meniscus of air-free water at a temperature of 68 °F (20 °C). The graduations on each tube shall be clearly numbered as shown in Figure C.1 and Figure C.2.

The tube graduation marks’ accuracy shall be volumetrically verified or gravimetrically certified before field use of the tube, in accordance with ASTM E 542, using National Institute of Standards and Technology (NIST)-traceable equipment. The verification or certification shall include a calibration check at each mark up through the 0.5 mL (1-part) mark; at the 1 mL, 1.5 mL, and 2 mL (2-part, 3-part, and 4-part) marks; and at the 50 mL and 100 mL (100-part and 200-part) marks. The tube shall not be used if the scale error at any mark exceeds the applicable tolerance from Tables 2 or 3. Tubes may have additional marks added at the 0.025 mL and 0.075 mL (0.05-part and 0.15-part) marks, which shall also be volumetrically verified or gravimetrically certified.

19

20	API Manual of Petroleum Measurement Standards 10.4

Figure C.1—8 in. (203 mm) Cone-shaped Centrifuge Tube

Manual of Petroleum Measurement Standards Chapter 10.4	21

Table C.1—Minimum Graduation Requirements and Maximum

Calibration Tolerances for 8 in. (203 mm) Cone-shaped Tubes

| Range (mL) | Subdivision (mL) | Volume Tolerance (mL) |
| --- | --- | --- |
| Above 0.0 to 0.1 | 0.05 | ±0.02 |
| Above 0.1 to 0.3 | 0.05 | ±0.03 |
| Above 0.3 to 0.5 | 0.05 | ±0.05 |
| Above 0.5 to 1.0 | 0.10 | ±0.05 |
| Above 1.0 to 2.0 | 0.10 | ±0.10 |
| Above 2.0 to 3.0 | 0.20 | ±0.10 |
| Above 3.0 to 5.0 | 0.50 | ±0.20 |
| Above 5.0 to 10.0 | 1.0 | ±0.50 |
| Above 10.0 to 25.0 | 5.0 | ±1.00 |
| Above 25.0 to 100.0 | 25.0 | ±1.00 |

Figure C.2—6 in. (167 mm) Cone-shaped Centrifuge Tube

22	API Manual of Petroleum Measurement Standards 10.4

Table C.2—Minimum Graduation Requirements and Maximum Calibration

Tolerances for 6 in. (167 mm) Cone-shaped Tubes

| Range (mL) | Subdivision (mL) | Volume Tolerance (mL) |
| --- | --- | --- |
| Above 0 to 0.1 | 0.05 | ±0.02 |
| Above 0.1 to 0.3 | 0.05 | ±0.03 |
| Above 0.3 to 0.5 | 0.05 | ±0.05 |
| Above 0.5 to1.0 | 0.10 | ±0.075 |
| Above 1.0 to1.5 | 0.10 | ±0.10 |
| Above 1.5 to 2.0 | 0.10 | ±0.20 |
| Above 2.0 to 3.0 | 0.20 | ±0.30 |
| Above 3.0 to 5.0 | 0.50 | ±0.50 |
| Above 5.0 to10.0 | 1.00 | ±0.75 |
| Above 10.0 to 25.0 | 5.00 | ±1.00 |
| Above 25.0 to 100.0 | Graduations at 50 and 100 | ±1.50 |

Annex D

(informative)

Abbreviated Procedure & Calculations—Determination of Sediment & Water

NOTE 1	An abbreviated procedure for determining sediment and water is provided below for convenience. Consult 9.2 earlier in this document for additional, alternative, or contingent requirements which may also apply.

NOTE 2	Users of instructions should not rely exclusively on the information contained in this document. Sound business,

scientific, engineering, and safety judgment should be used in employing the information contained herein.

NOTE 3	Where applicable, authorities having jurisdiction should be consulted.

NOTE 4 Work sites and equipment operations may differ. Users are solely responsible for assessing their specific equipment and premises in determining the appropriateness of applying the instructions. At all times users should employ sound business, scientific, engineering, and judgement safety when using this standard.

23

Annex E

(informative)

Abbreviated Procedure & Calculations—Determination of Sediment Only

NOTE 1	An abbreviated procedure for determining sediment and water is provided below for convenience. Consult 9.2 earlier in this document for additional, alternative, or contingent requirements which may also apply.

NOTE 2	Users of instructions should not rely exclusively on the information contained in this document. Sound business,

scientific, engineering, and safety judgment should be used in employing the information contained herein.

NOTE 3	Where applicable, authorities having jurisdiction should be consulted.

NOTE 4 Work sites and equipment operations may differ. Users are solely responsible for assessing their specific equipment and premises in determining the appropriateness of applying the instructions. At all times users should employ sound business, scientific, engineering, and judgement safety when using this standard.

24

# Bibliography

- API MPMS Chapter 8.4, Standard Practice for Sampling and Handling Fuels for Volatility Measurement

- U.S. DOL Title 29 CFR 1910,2 Subpart Z—Toxic and Hazardous Substances

- ACGIH,3 “Threshold Limit Values and Biological Exposure Indices for 1987–1988,” Cincinnati, Ohio, 1987

2 U.S. Department of Labor, Occupational Safety and Health Administration (OSHA), 200 Constitution Avenue, NW, Washington, D.C. 20210, www.osha.gov.

3  ACGIH, 1330 Kemper Meadow Drive, Cincinnati, Ohio 45240, www.acgih.org.

25

200 Massachusetts Avenue, NW Suite 1100

Washington, DC 20001-5571 USA

202-682-8000

Additional copies are available online at www.api.org/pubs

Phone Orders:	1-800-854-7179 (Toll-free in the U.S. and Canada) 303-397-7956	(Local and International)

Fax Orders:	303-397-2740

Information about API publications, programs and services is available on the web at www.api.org.

Product No. H100405